## Slide 1: CLAS12 Low Threshold Cherenkov Counter
Assessment of the LTCC Project
Group Leader: Maurizio Ungaro
Tech Leader: David Anderson
Eng. Leader: Bob Miller
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope
Work Completed
  Box Cut, Nose, Mirrors re-location
  Back-wall, Connectors
  Mirror Coating
  Winston Cones Coating
  PMT coating
  PMT Divider
  Window fabrication
  Mirror Support
  Mirrors Alignment
  Calibration
Work Remaining
  Procurement
  Labor
  Commissioning
Turbo
  How to meet FY15

## Slide 2: Outline
LTCC Overview
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 3: LTCC Overview
DC
FTOF
PCAL, EC
HTCC
Central Detector
BST, BMT, FMT, CTOF, CND
LTCC
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope

## Slide 4: LTCC Overview
108 lightweight mirrors
36 Winston Cones
36 5’’ Photonis X4500B PMT
36 Magnetic Shields
C4F10 Gas, r.i. 1.00134
CLAS6: e-/π discrimination
π threshold: 2.6 GeV
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope
One
Sector

## Slide 5: LTCC CLAS6 Performance
lab ɸ
lab θ
17 years of “all on” PMTs (last 2 years just 1 dead)
good electron ID:
most: 10-20 nphe
some region: 5-10 nphe
PMT gain low (x10 multiplier was necessary)
Lower efficiency in the mid-sector
$50K / year in gas leaks
Mirror support malfunction
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope
LTCC was used for e-/π discrimination

## Slide 6: LTCC: from CLAS6 to CLAS12
HTCC will do
e-/π discrimination
LTCC could do
π/K discrimination
Fit between DC and FTOF: box is too large
Decrease inefficient regions
Increase PMT gain, split output (FADC, TDC)
Minimize gas leaks
Better mirrors support
Re-fabricate windows
Increase light yield for π
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope
Refurbishment
Challenges:

## Slide 7: 12CBLT Project Overview
Re-coat all mirrors
Re-coat WCs
Wavelength shifting coating of the PMTs
Add “nose” to increase gas volume
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope
Fit between DC and FTOF: box is too large
Cut Sides
Relocate segments 16-17-18
Redesign, replace back-wall frame
2. Decrease inefficient regions
Add mirror pieces to cover gaps
3. Increase PMT gain, split output (FADC, TDC)
Vladimir Popov base modification: 2 x10 output
4. Minimize gas leaks
Use Hermetical signals, HV connectors
Seal the box from the inside
Window: glue + sealant
5. Better mirrors support
Design new spine
6. Re-fabricate windows
Lamination (company)
Seaming (in house)
7. Increase light yield for π

## Slide 8: Photo-electrons yield
1.5m of C4F10 (with its transparency)
Mirrors, WC reflectivity: original, refurbished
PMT type: original, refurbished
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope

## Slide 9: Photo-electrons yield
1.5m of C4F10 (with its transparency)
Mirrors, WC reflectivity: original, refurbished
PMT type: original, refurbished
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope

## Slide 10: Photo-electrons yield
1.5m of C4F10 (with its transparency)
Mirrors, WC reflectivity: original, refurbished
PMT type: original, refurbished
LTCC Overview
  Technical Parameters
  CLAS6 Performance
  Physics Motivation
  Project Scope
CLAS12 light yield / CLAS light yield
Purpose: obtain same light yield (for pions) as we had for electrons

## Slide 11: Outline
Work Completed
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 12: Box Cut, Nose, Mirrors relocation
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 13: Box Cut, Nose, Mirrors relocation
  Overview
  Work Completed
  Work Remaining
  Turbo
x
x

## Slide 14: Back-wall, Connectors
Back-wall re-designed:
New frame
Old back-wall fitted on new frame, sealed
New hermetical connectors
One more PMT output
Back-wall: COMPLETED
Frame: $45,436.00
Connectors: $7,566.50
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 15: Mirrors
2 reflections only: 30% of the time
3 reflections ~40% of the time
4 reflections ~30% of the time
3 reflections visible:
~27.5% (original, 65% 1 reflection)
~68.1% (recoated, 88% 1 reflection)
3 reflections UV:
~9.1% (original, 45% 1 reflection)
~31.3% (recoated, 70% 1 reflection)
X ~2.4 better
X ~3.7 better
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 16: Cylindrical Mirrors
36 / box, 216 total
3 companies considered, tests from all
Awarded to ECI
Process: mirror in vacuum chamber, Al + MgF2 evaporation
Project Complete. Cost: $11,942.50
  Overview
  Work Completed
  Work Remaining
  Turbo
Cylindrical Mirrors Coating: COMPLETED

## Slide 17: Elliptical, Hyperbolic Mirrors
Elliptical Mirrors: 36 / box, 216 total
Hyperbolic Mirrors: 36 / box, 216 total
Problem: cannot be re-coated:
Mirror outgas in vacuum chamber.
Many mirrors do not fit vacuum chambers.
All 3 companies failed on re-coating.
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 18: Elliptical, Hyperbolic Mirrors
Elliptical Mirrors: 36 / box, 216 total
Hyperbolic Mirrors: 36 / box, 216 total
Problem: cannot be re-coated:
Solution:
Coat lexan strips
Glue lexan strips on top of mirror
Challenge:
Protect lexan strips while gluing to mirror
Final Product:
Only ECI could deliver lexan with layer of protection that did not degrade reflectivity
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 19: Elliptical, Hyperbolic Mirrors
Clean room environment
Spray glue pumped outside
Mirrors vaccum-holded
Lexan Strips rolled on mirror
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 20: Elliptical, Hyperbolic Mirrors
Project Complete
Materials Cost: $53,100.00
Item Quantity Part Description USD Price
A 190 9" x 36" x 0.010” coated sheet $ 133.65 each
B 150 10" x 36" x 0.010” coated sheet $ 195.10 each
C 6 10" x 36" x 0.050” coated sheet $ 368.90 each
B
E
F
O
R
E
A
F
T
E
R
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 21: Winston Cones Coating
• Small:
  o Height: 18cm
  o Parallel Plate distance: 14cm
  o Radius at the top: 20cm
  o Radius at the bottom: 11cm
  o Material: copper (electro-formed)
• Medium:
  o Height: 22cm
  o Parallel Plate distance: 15cm
  o Radius at the top: 20cm
  o Radius at the bottom: 11cm
  o Material: 0.2” plastic (vacuum pressed)
• Large:
  o Height: 30cm
  o Parallel Plate distance: 18cm
  o Radius at the top: 22cm
  o Radius at the bottom: 11cm
  o Material: copper (electro-formed)
ECI Awarded: $500 / apiece
Budget: ~80K, or 160 cones out of 216.
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 22: Winston Cones Coating
Original WCs Reflectivity UV Average
DSG Group Performed Tests
Setup:
Monochromator, splitter
Incident beam ~10-20 degrees
Reference photodiode
Labview to control wavelength
Labview to read two nanoAmmeters
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 23: Winston Cones Coating
Original WCs Reflectivity UV Average
  Overview
  Work Completed
  Work Remaining
  Turbo
Sector 1 |  | Sector 2 |  | Sector 3 |  | Sector 4 |  | Sector 5 |  | Sector 6 | 
Left | Right | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right
0.67 | 0.30 | 0.67 | 0.70 | 0.56 | 0.56 | 0.51 | 0.81 | 0.40 | 0.44 | 0.72 | 0.80
0.26 | 0.36 | 0.62 | 0.59 | 0.64 | 0.63 | 0.86 | 0.65 | 0.37 | 0.37 | 0.77 | 0.38
0.70 | 0.45 | 0.64 | 0.65 | 0.38 | 0.71 | 0.42 | 0.51 | 0.33 | 0.39 | 0.50 | 0.47
0.33 | 0.27 | 0.79 | 0.76 | 0.57 | 0.76 | 0.75 | 0.64 | 0.32 | 0.81 | 0.46 | 0.41
0.32 | 0.27 | 0.70 | 0.72 | 0.71 | 0.51 | 0.73 | 0.69 | 0.45 | 0.65 | 0.89 | 0.24
0.57 | 0.45 | 0.69 | 0.65 | 0.59 | 0.61 | 0.55 | 0.78 | 0.33 | 0.42 | 0.44 | 0.57
0.31 | 0.49 | 0.68 | 0.65 | 0.71 | 0.81 | 0.75 | 0.83 | 0.40 | 0.41 | 0.37 | 0.54
0.33 | 0.459029627 | 0.53 | 0.56 | 0.76 | 0.76 | 0.78 | 0.74 | 0.43 | 0.50 | 0.52 | 0.32
0.69 | 0.54 | 0.47 | 0.78 | 0.67 | 0.75 | 0.74 | 0.72 | 0.45 | 0.47 | 0.55 | 0.47
0.33 | 0.30 | 0.68 | 0.71 | 0.71 | 0.75 | 0.68 | 0.83 | 0.08 | 0.39 | 0.42 | 0.44
0.78 | 0.59 | 0.79 | 0.80 | 0.77 | 0.76 | 0.64 | 0.74 | 0.77 | 0.68 | 0.77 | 0.75
0.77 | 0.72 | 0.56 | 0.77 | 0.76 | 0.71 | 0.73 | 0.74 | 0.60 | 0.90 | 0.75 | 0.80
0.22 | 0.73 | 0.72 | 0.69 | 0.46 | 0.48 | 0.72 | 0.68 | 0.69 | 0.75 | 0.50 | 0.64
0.51 | 0.59 | 0.59 | 0.44 | 0.43 | 0.53 | 0.63 | 0.46 | 0.62 | 0.68 | 0.44 | 0.66
0.47 | 0.70 | 0.49 | 0.71 | 0.63 | 0.48 | 0.82 | 0.64 | 0.65 | 0.70 | 0.67 | 0.42
0.72 | 0.66 | 0.72 | 0.77 | 0.75 | 0.66 | 0.71 | 0.52 | 0.70 | 0.70 | 0.57 | 0.62
0.66 | 0.74 | 0.78 | 0.43 | 0.59 | 0.64 | 0.52 | 0.71 | 0.54 | 0.69 | 0.19 | 0.66
0.77 | 0.54 | 0.78 | 0.74 | 0.76 | 0.63 | 0.63 | 0.66 | 0.69 | 0.55 | 0.69 | 0.59
BAD
GOOD
10 passed quality tests
Criteria:
R >= 0.8

## Slide 24: Winston Cones Coating
Coated WCs Reflectivity UV Average
  Overview
  Work Completed
  Work Remaining
  Turbo
Sector 1 |  | Sector 2 |  | Sector 3 |  | Sector 4 |  | Sector 5 |  | Sector 6 | 
Left | Right | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right
0.67 | 0.81 | 0.67 | 0.65 | 0.77 | 0.56 | 0.51 | 0.81 | 0.86 | 0.81 | 0.81 | 0.80
0.74 | 0.83 | 0.85 | 0.59 | 0.64 | 0.69 | 0.86 | 0.65 | 0.85 | 0.85 | 0.77 | 0.82
0.70 | 0.84 | 0.64 | 0.65 | 0.83 | 0.71 | 0.85 | 0.73 | 0.73 | 0.81 | 0.82 | 0.79
0.75 | 0.78 | 0.79 | 0.76 | 0.57 | 0.76 | 0.84 | 0.77 | 0.54 | 0.81 | 0.80 | 0.84
0.81 | 0.80 | 0.71 | 0.82 | 0.81 | 0.84 | 0.75 | 0.69 | 0.86 | 0.80 | 0.89 | 0.81
0.86 | 0.83 | 0.69 | 0.65 | 0.81 | 0.81 | 0.78 | 0.78 | 0.82 | 0.84 | 0.79 | 0.57
0.80 | 0.86 | 0.68 | 0.85 | 0.81 | 0.75 | 0.69 | 0.83 | 0.81 | 0.81 | 0.83 | 0.79
0.87 | 0.459029627 | 0.75 | 0.80 | 0.74 | 0.76 | 0.78 | 0.78 | 0.84 | 0.81 | 0.85 | 0.82
0.69 | 0.80 | 0.84 | 0.78 | 0.67 | 0.75 | 0.76 | 0.72 | 0.82 | 0.84 | 0.84 | 0.86
0.83 | 0.30 | 0.82 | 0.71 | 0.81 | 0.75 | 0.68 | 0.83 | 0.08 | 0.78 | 0.80 | 0.81
0.78 | 0.59 | 0.79 | 0.80 | 0.77 | 0.76 | 0.64 | 0.66 | 0.77 | 0.56 | 0.77 | 0.75
0.77 | 0.68 | 0.56 | 0.77 | 0.52 | 0.71 | 0.65 | 0.68 | 0.60 | 0.90 | 0.65 | 0.80
0.55 | 0.82 | 0.72 | 0.85 | 0.53 | 0.50 | 0.89 | 0.68 | 0.70 | 0.75 | 0.50 | 0.75
0.51 | 0.66 | 0.70 | 0.44 | 0.48 | 0.55 | 0.63 | 0.46 | 0.62 | 0.67 | 0.44 | 0.66
0.77 | 0.75 | 0.48 | 0.73 | 0.63 | 0.48 | 0.82 | 0.64 | 0.65 | 0.68 | 0.67 | 0.42
0.86 | 0.91 | 0.81 | 0.77 | 0.82 | 0.81 | 0.69 | 0.52 | 0.89 | 0.67 | 0.45 | 0.62
0.73 | 0.82 | 0.78 | 0.53 | 0.78 | 0.72 | 0.52 | 0.75 | 0.54 | 0.76 | 0.80 | 0.66
0.77 | 0.58 | 0.78 | 0.86 | 0.68 | 0.68 | 0.55 | 0.79 | 0.73 | 0.55 | 0.62 | 0.84
75 passed our quality tests, 117 if criteria is R >= 0.75
BAD
GOOD
Criteria:
R >= 0.8

## Slide 25: Winston Cones Coating
Issues:
Many plastic and most large cones show same or worse reflectivity.
  Problem #1: previous lamination was flaky.
  Solution: We removed + polished
  The flaky part on and sent back to ECI for re-coat.
  The test back looked good so we went into production:
70 cones were delaminated in-house.
  Out of the 70 sent back, only ~20 show definite improvements.
  10 are so-so and the remaining show degradation.
The problem is the previous lamination prevents a good deposition
In progress:
Researching de-lamination by chemical means
Researching de-lamination by rough polishing (one is completed and ready to test)
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 26: Winston Cones Coating
WC STATUS:
36 installed (first box)
89 ready to install
21 cones at ECI
70 that would need de-lamination
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 27: PMT Coating
p-terphenyl PMT coating proved successful in many experiments (at JLab as well)
Improve quantum efficiency of UV-glass PMTs to rival that of quartz PMTs
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 28: PMT Coating Procedure
Vacuum (10-6) ~a few hours
Clean PMT (acetone, alcohol)
Load p-terphenyl boat (~1.25g)
Heat PT boat, 1500C
Shutter close PMTs
Switch Boat
Heat MgF2 (12000C)
Open Shutter for fast MgF2 deposition
2 LED (265nm and 290nm) at Temple for QA
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 29: PMT Coating Measurement
2 Coated PMTs
1 Reference PMT
SPE matched
Rates Matched at 400 nm
(~10 HV change from SPE match)
Lamp Intensity so that:
Rates: ~100 KHz
Wavelength range:
200-400nm
Switched PMT / Reference PMTs positions: <10% effect
Shown: gain average between the 2 positions
JLab Monochromator: 200-400nm
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 30: PMT Coating Measurement
Points: ALL coated PMTs LED
Gain Factor average
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 31: PMT Coating
Project Complete
200 PMTs, $330 apiece
Cost: $66,000.00
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 32: PMT Divider
Vladimir Popov Design (Photonis wanted it!)
Provides 2 identical signal: X10 amplification (no penalties)
240 pcp boards populated by Flextron Circuit Assembly
167 divider modifications at JLab
  Overview
  Work Completed
  Work Remaining
  Turbo
CLAS: PMT signal was divided and amplified by UVA electronics modules no longer available.
Popov’s modification can be used instead

## Slide 33: PMT Divider
self trigger (10-100 mV)
2 signals are identical
Voltage: 100V LESS than w/o modifier!
Data acquisition (raw mode) to test base AND pmt
  Overview
  Work Completed
  Work Remaining
  Turbo
integration region
raw FADC Spectrum

## Slide 34: PMT Divider
FADC Pulse Spectrum: SPE loud & clear!
Project cost:
~$5500 (materials)
6 FTE weeks DSG
3 days Popov’s Time
V.U. will complete mod
  Overview
  Work Completed
  Work Remaining
  Turbo
Project: COMPLETED

## Slide 35: Mirrors Support
Aluminum Honeycomb “spine”. Attached with strips, plates and glue
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 36: Mirrors Support
New Spine:
1” diameter, 1mm thick carbon fiber
22 mils stainless steel cabling to the box sides
Adjustment buckles, springs
Spine is “free to move” inside tube bolted on wall, 1/2” free space to allow for the 8mm movement of the back-wall
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 37: Mirrors Support Test
  Overview
  Work Completed
  Work Remaining
  Turbo
Box is vertical
Box is horizontal

## Slide 38: Mirrors Alignment
Laser line (200) at the surveyed target position
Laser line (30) on frame near upstream window (virtual pion) give same results
Optimize Elliptical Mirror alignment
Optimize Hyperbolic Mirror alignment
Fix to mirrors support
Optimize both mirrors to collect as much light as possible
  Overview
  Work Completed
  Work Remaining
  Turbo
WC missing from simulation

## Slide 39: Mirrors Alignment
Alignment is completed on the first box
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 40: Window Fabrication
Need: 15’x15’ upstream and downstream windows, for HTCC and LTCC
62” roll of black tedlar, 200 lbs (1300 yards) : $12,600 (HTCC, DuPont Tedlar)
Lamination: $20,000 (LTCC, Madico)
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 41: Window Tests
Weight(lbs) | Length(in.)
0 | 10.31
50 | 10.44
100 | 10.50
150 | 10.56
200 | 10.63
250 | 10.75
300 | 11.06
388 | Failed at seam
Seam
Lamination
Force (lbs) | Length | Stretch (in) | Stress (psi)
6 | 9.1250 | 0.0000 | 250.0
37 | 9.1250 | 0.0000 | 1541.7
76 | 9.1250 | 0.0000 | 3166.7
117 | 9.1875 | 0.0625 | 4875.0
160 | 9.2500 | 0.1250 | 6666.7
200 | 9.4375 | 0.3125 | 8333.3
240 | 9.8125 | 0.6875 | 10000.0
280 | 10.5000 | 1.3750 | 11666.7
440 | 12.1250 | 3.0000 | 18333.3
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 42: Window On LTCC
Downstream window sealant curing
Attachment to box:
Sealant: silicon
Glue: G-Flex
2 layers of silicon sandwich a 4” wide band of g-flex
  Overview
  Work Completed
  Work Remaining
  Turbo
Window assembly project:
COMPLETED
All windows + 4 spares

## Slide 43: LTCC Calibrations
CLAS12 Calibrations
HTCC and LTCC
HV Matching
SPE analysis
CCDB I/O
CAEN SY4527 HV I/O
ROOT Output
PMT, bases currently being tested using the a beta version of the software
Missing: Reading from ET ring
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 44: Outline
Work Remaining
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 45: LTCC Remaining Procurement
Winston Cones: $36,868, ~70 cones
Delamination research in progress:
by chemical means ($30-60 apiece, at Nicoform)
  Cost: $4200 – Estimate by Nicoform (company that produced the HTCC WC)
  Budget Increase
by polishing (one is completed and ready to test)
V.U. Labor
No other remaining procurement
Work Remaining
  Procurement
  Labor
  Commissioning

## Slide 46: LTCC Remaining Labor
First box being closed up as we speak:
Back-wall installation
Hardware population
Mirror Alignment and support
Window installation
done for first box, need to do it for the remaining 5 boxes.
All procedures checked out.
Issue: alignment, window activities cannot be done at the same time as hardware population
Estimate: 4-5 weeks / box (2 Techs) = 50 Mweeks
1 Tech from Hall-D
1 Visiting User
Work Remaining
  Procurement
  Labor
  Commissioning

## Slide 47: LTCC Remaining Activities
Activity ID | Name | Activity Start | Activity Finish | Remaining Budgeted MWeeks or Proc. Cost
242242512, 517, 525, 530 | Winston Cone and PMT work | 1/5/15 | 9/21/15 | $36,868 + 2.6
242242550-575 | Complete Box 1 | 1/23/15 | 6/30/15 | 3.6
242242590-620 | Complete Box 2 | 1/23/15 | 8/14/15 | 9.5
242242635-665 | Complete Box 3 | 2/19/15 | 9/22/15 | 9.8
242242680-710 | Complete Box 4 | 2/23/15 | 10/29/15 | 9.7
242242725-755 | Complete Box 5 | 2/16/15 | 12/9/15 | 11.0
242242770-805 | Complete Box 6 | 2/16/15 | 1/27/16 | 11.3
24267175, 180, 190, 195, 215 | LTCC Installation Tooling and Mounting | 8/4/15 | 2/19/16 | $5,847 + 15
2822X045, 050 | Calibration Software | 1/5/15 | 6/30/15 | 3
2822X065, 070 | Survey/Cable/Test/Repair | 2/22/16 | 3/2/16 | 5.4
~ 60 Mweeks for LTCC
Work Remaining
  Procurement
  Labor
  Commissioning

## Slide 48: LTCC Commissioning
C4F10 Gas load test
PMTs commissioning:
  All channels powered, connected to DAQ.
  Calibration suite
SPE Calibration:
  HV Matching to align SPE peaks
Slow Control:
  Straightforward, HV control using CLAS12 tools
Work Remaining
  Procurement
  Labor
  Commissioning
OFF PROJECT
Monitoring:
  Straightforward, using CLAS12 tools
Beam ON additional commissioning:
  Timing Calibration (Calibration Suite)
  HV Matching in torus field
1 year: Oct 2015 to Oct 2016
20% FTE Scientist
M. Ungaro

## Slide 49: Turbo
  How to meet FY15
  Overview
  Work Completed
  Work Remaining
  Turbo

## Slide 50: Turbo
  Overview
  Work Completed
  Work Remaining
  Turbo
Second production line, in the nearby HTCC clean tent
Alignment can be done in parallel with hardware and window installation (including leak checks).
Nearby location is critical for productivity (other locations would not be worth it)
To meet FY15
No MATERIAL COST or MANPOWER increase
HTCC clean room availability ASAP is critical

## Slide 51: Summary
All hardware component are in place, except WC.
First box is completed
One remaining procurement: WC. $36,868
  Possible budget increase: $4200
  Research in progress to solve the related problems.
Remaining Labor:
  10 MT MWeeks / box = 50 MT Mweeks
  1.5 Scient. Mweeks / Box = 7.5 Mweeks
  5 Mweeks VU
Second production line necessary to meet FY 2015
Detector Installation scheduled for Spring 2016

## Slide 52: Backup
Sector 1 |  | Sector 2 |  | Sector 3 |  | Sector 4 |  | Sector 5 |  | Sector 6 | 
Left | Right | Left | Right | Left | Right | Left | Right | Left | Right | Left | Right
0.67 | 0.81 | 0.67 | 0.65 | 0.77 | 0.56 | 0.51 | 0.81 | 0.86 | 0.81 | 0.81 | 0.80
0.74 | 0.83 | 0.85 | 0.59 | 0.64 | 0.69 | 0.86 | 0.65 | 0.85 | 0.85 | 0.77 | 0.82
0.70 | 0.84 | 0.64 | 0.65 | 0.83 | 0.71 | 0.85 | 0.73 | 0.73 | 0.81 | 0.82 | 0.79
0.75 | 0.78 | 0.79 | 0.76 | 0.57 | 0.76 | 0.84 | 0.77 | 0.54 | 0.81 | 0.80 | 0.84
0.81 | 0.80 | 0.71 | 0.82 | 0.81 | 0.84 | 0.75 | 0.69 | 0.86 | 0.80 | 0.89 | 0.81
0.86 | 0.83 | 0.69 | 0.65 | 0.81 | 0.81 | 0.78 | 0.78 | 0.82 | 0.84 | 0.79 | 0.57
0.80 | 0.86 | 0.68 | 0.85 | 0.81 | 0.75 | 0.69 | 0.83 | 0.81 | 0.81 | 0.83 | 0.79
0.87 | 0.459029627 | 0.75 | 0.80 | 0.74 | 0.76 | 0.78 | 0.78 | 0.84 | 0.81 | 0.85 | 0.82
0.69 | 0.80 | 0.84 | 0.78 | 0.67 | 0.75 | 0.76 | 0.72 | 0.82 | 0.84 | 0.84 | 0.86
0.83 | 0.30 | 0.82 | 0.71 | 0.81 | 0.75 | 0.68 | 0.83 | 0.08 | 0.78 | 0.80 | 0.81
0.78 | 0.59 | 0.79 | 0.80 | 0.77 | 0.76 | 0.64 | 0.66 | 0.77 | 0.56 | 0.77 | 0.75
0.77 | 0.68 | 0.56 | 0.77 | 0.52 | 0.71 | 0.65 | 0.68 | 0.60 | 0.90 | 0.65 | 0.80
0.55 | 0.82 | 0.72 | 0.85 | 0.53 | 0.50 | 0.89 | 0.68 | 0.70 | 0.75 | 0.50 | 0.75
0.51 | 0.66 | 0.70 | 0.44 | 0.48 | 0.55 | 0.63 | 0.46 | 0.62 | 0.67 | 0.44 | 0.66
0.77 | 0.75 | 0.48 | 0.73 | 0.63 | 0.48 | 0.82 | 0.64 | 0.65 | 0.68 | 0.67 | 0.42
0.86 | 0.91 | 0.81 | 0.77 | 0.82 | 0.81 | 0.69 | 0.52 | 0.89 | 0.67 | 0.45 | 0.62
0.73 | 0.82 | 0.78 | 0.53 | 0.78 | 0.72 | 0.52 | 0.75 | 0.54 | 0.76 | 0.80 | 0.66
0.77 | 0.58 | 0.78 | 0.86 | 0.68 | 0.68 | 0.55 | 0.79 | 0.73 | 0.55 | 0.62 | 0.84
117 good
BAD
GOOD
Criteria:
R >= 0.75