## Slide 1: CLAS12 Central Time-of-Flight
CTOF Detector Assessment
Group Leader:
Daniel S. Carman
Outline:
  Overview
  Project Status
  Remaining Tasks
    On Project
    Off Project

## Slide 2
Overview

## Slide 3: CLAS12
Base equipment
Forward Detector (FD)
  - Torus magnet
  HT Cherenkov Counter
  Drift Chamber system
  LT Cherenkov Counter
  Forward Time-of-Flight
  Pre-shower Calorimeter
  E.M. Calorimeter
Central Detector (CD)
  Solenoid magnet
  Barrel Silicon Tracker
  Central Time-of-Flight
Beamline
  Targets
  Moller Polarimeter
  Photon Tagger
Upgrade to base equipment
- MicroMegas
  Central Neutron Detector
  Forward Tagger
  RICH detector (1 sector)
  Polarized target (long.)
MM
CND
http://www.jlab.org/Hall-B/clas12-web/

## Slide 4: System Overview
PARAMETER | DESIGN VALUE | #
CTOF Barrel | 48 counters; double-sided readout; R=25 cm | 
Angular Coverage | q: (35o , 125o) f: (-180o,180o) | 
Scint. Dimensions | BC-408 Trap. cross section ~3 x 3 x 92 cm3 | 48
PMTs | R2083 (PC 46 mm, H2431-MOD assembly) | 96
Upstream LGs | OD=2”, 1m-long, focusing design, straight | 48
Downstream LGs | OD=2”,1.6 m-long, focusing design, bent 135o | 48
Magnetic Shields | Triple layer Co-NETIC/Hi-Perm 49/steel with compensating coil around inner layer | 96
Design Resolution | 60 ps | 
PID: |  | 
p/K Separation | to 0.64 GeV (3.3s) | 
K/p Separation | to 1.00 GeV (3.3s) | 
p/p Separation | to 1.25 GeV (3.3s) | 
Operating Environment:
CTOF installed in CLAS12 5-T superconducting solenoid
Fringe field at CTOF PMT location at level of 0.5 - 1 kG
CND outside; SVT inside
CTOF Design:
low-pitch angle
high-pitch angle

## Slide 5: CTOF PID Specifications
PID: | 3.3s Separation
p/K | 0.64 GeV/c
K/p | 1.00 GeV/c
p/p | 1.25 GeV/c
Nominal CTOF average timing specification is 60 ps/counter
CTOF is the primary PID detector in the CLAS12 Central Detector

## Slide 6
Project Status
Counter Assembly
Cosmic Ray Testing
Software Development
Remaining Design/Engineering

## Slide 7: Counter Assembly
Counter Assembly Status:
All 48 counter elements have been polished and then glued together using precision assembly table
Wrapping of all counters complete
VM2002 reflective layer, Tedlar outer wrap
Assembly Table
- Installed in custom storage racks
- Magnetic shields attached
PMTs installed and oriented
Light leak checking complete
Counter Storage Racks
Work Complete

## Slide 8: Counter Wrapping
Cold-formed Tedlar using new tooling
Careful attention to material layers given tight counter-to-counter tolerances
  Counter-to-counter spacing < 1 mm
Problems caused 2 month schedule delay
In Jan. 2015 it was discovered that counter Tedlar wrapping had to be redone
New techniques and tooling were developed and tested
Re-wrapping and light-leak testing completed in Mar. 2015
Work Complete

## Slide 9
Stabilized Dividers
PMT/Voltage Divider Assemblies
PMT Gain vs. Current/Rate
Pulser Peak Position vs. Current
Blue points – original Hamamatsu PMT/divider (H2431)
Nominal current limit spec. at 100 mA
Red points – rate-stabilized divider (V. Popov)
Stable timing performance up to 200 mA
mezzanine board
Work Complete

## Slide 10: CTOF PMTs in Solenoid Fringe Field
Upstream low-pitch PMT: (z,y)1 = (493 G, 956 G), (z,y)2 = (193 G, 342 G)
Upstream high-pitch PMT: (z,y)1 = (192 G, 981 G), (z,y)2 = (78 G, 337 G)
Downstream PMTs in twice lower field
Renuka Rajput-Ghoshal
1
2

## Slide 11: Dynamic Magnetic Shields
Shield design verified by 2D (Poisson) and 3D (Opera) modeling
< 1 G at photocathode
coordinate x
- 3 passive layers
  Co-NETIC (inner)
  Hi-Perm 49 (middle)
  Steel 1008 (outer)
1 active layer
  2 compensation coils wound on inner layer for field shaping
V. Baturin et al., Nucl. Inst. Meth. A 664, 11 (2011)
shield design V. Baturin

## Slide 12: Shield Parts – “Exploded” View
steel
Hi-perm
Co-NETIC

## Slide 13: Magnetic Shield Assembly
Bench assembly
Shield weight = 40 lbs
Work Complete

## Slide 14: CTOF Compensation Coils
Each coil can produce a cancelling field of up to ~50 G
Work Complete

## Slide 15: Magnetic Shield Testing
Study final field reduction factor of CTOF production shields:
  Use FROST 5-T superconducting solenoid – fringe field
  Position shield in the same field/gradient as in CLAS12 solenoid field
  Measure internal remnant field vs. operating conditions
Shield Test Plan:
Study field reduction factor relative to design value
Study performance limits of shield design
Understand limits of ferromagnetic saturation
Optimize compensation coil configuration
Study impact of remnant field on PMT pulse shape and azimuthal orientation
Shield tests completed June 8-12

## Slide 16: CTOF Magnetic Shield Testing
Transverse coordinate (cm)
CTOF PMTs positioned in fields up to 1.1 kG
Internal fields at PMT < 0.7 G in fields that matched those expected in CLAS12 Solenoid fringe field
Shield tested in fields up to Bmax = 2.2 kG with good performance
ph-cath < 0.7 G
mid-dyn < 0.2 G
1st dyn < 0.1 G
(G)
3 passive layers + 1 active layer
Work Complete

## Slide 17: CTOF Technical Notes
Design and Performance Tests of the CTOF Magnetic Shields
CTOF Geometry Document
CTOF Calibration Constants
CTOF Counter Wrapping Procedures
CTOF Magnetic Shield Attachment Procedures
LMS Design Considerations
TOF Monte Carlo Simulation Details
2015 Monthly Progress Reports
See CLAS12 CTOF web page at:
http://www.jlab.org/Hall-B/ctof
CTOF Documentation
  CLAS-Notes, CLAS12-Notes, Internal Documents, Reports
Photographs of all things CTOF

## Slide 18
Project Status
Counter Assembly
Cosmic Ray Testing
Software Development
Remaining Design/Engineering

## Slide 19: CTOF Cosmic Ray Studies
Counters now connected to DAQ system for cosmic ray studies
gain matching
calibration
time resolution optimization
DAQ setup with production readout:
CFDs (Ortec 935)
High-res TDCs (CAEN 1290A)
FADCs (JLab modules)
CODA readout – EVIO format
CTOF DAQ Test Stand

## Slide 20
CTOF HV Cable Connections
CTOF Signal Cable Connections
PMT SHV M
SHV F SHV F
Cable RG-59
(18’ up, 23’ dn)
HV Patch
SHV M
PMT BNC F
Anode
BNC M Lemo M
Cable EML-195
Flat Lemo M
Cable RG-316
splitter panel
up: 52’, 53.6’, 55.2’
dn: 68’, 69.6’, 71.2’
(5’)
Lemo M Lemo M
Cable RG-316
(5’)
disc
FADC
Other cables: Lemo M Lemo M (disc to TDC); Lemo M Lemo M (CFD delay)
RG-316 (4’) RG-174 (2 ns)
CTOF Cabling Scheme
Work Complete

## Slide 21: PMT Gain Matching
CTOF counter PMT gain matching complete
Verify PMT settings and eliminate mis-cabling

## Slide 22: Attenuation Lengths
Attenuation Length l Data

## Slide 23: Time Resolution Studies
Compute variance:
Setup triplet of assembled CTOF counters:
(ns)
(ns)
(cm)

## Slide 24: Time Resolution Studies
Preliminary

## Slide 25
Project Status
Counter Assembly
Cosmic Ray Testing
Software Development
Remaining Design/Engineering

## Slide 26: CTOF Slow Controls
CLAS12 HVPS Slow Controls EPICS interface developed by V. Sytnik and K. Livingston (contributed University labor)
Work Complete

## Slide 27: CTOF Monitoring
Trigger scalers
tools now under development – off project
Sergey Boyarinov

## Slide 28
CTOF Calibration Suite
The CTOF calibration suite will be developed directly from the code suite being developed for the FTOF system.
ROOT/c++ implementation:
  H. Lu (Univ. of Iowa)
JAVA implementation:
  L. Clark (Univ. of Glasgow)
GUI development:
  L. Clark (Univ. of Glasgow)
Dedicated development of CTOF suite will begin later this year
Contributed University Labor – Off Project

## Slide 29
Project Status
Counter Assembly
Cosmic Ray Testing
Software Development
Remaining Design/Engineering

## Slide 30: Light Monitoring System
Status:
Design of system nearly complete
  (A/B) – light source/driver developed at JLab
  (C) – integrating sphere commercial unit
  (D) – fiber bundle assembly commercial unit
  (E) – fiber mounting blocks from local shop
Block Diagram
Fiber Bundle Assembly
LED driver
A
B
C
D
E
Similar system already installed and tested for HTCC
Counter time resolution/gain monitoring system; prescaled part of datastream
5-m long
- Procurement to start in July 2015

## Slide 31: Support Structure
Status: Main design elements completed; drawing set sent out for RFI in early May
Designed complicated by presence of CND that is installed first in solenoid and weight of magnetic shields
PR submission to begin this week
upstream support
downstream support

## Slide 32: Installation Scheme & Tooling
Step 1: Mount counter to installation tooling
Step 2: Move counter into Solenoid; rotate in azimuth
Step 3: Rotate counter toward Solenoid bore
Step 4: Attach the counter to the support structure
Status: Design (65%) and engineering analysis (33%) in progress (June reporting)

## Slide 33
CTOF Installation
Sequence
Cyril Wiggins

## Slide 34
Install ANCHOR BRACKETS onto UPSTREAM INTERFACE RING of SOLENOID

## Slide 35
Install CENTRAL NEUTRON DETECTOR onto UPSTREAM INTERFACE RING

## Slide 37
Install CTOF support structure onto anchor brackets

## Slide 38
Install counters #1 & #2;
- Specially assembled, as the first of the array
- Require special installation because of close proximity to the solenoid cart

## Slide 39
Install first quadrant of counters from the bottom up

## Slide 40
Install second quadrant of counters

## Slide 41
Install third and fourth quadrants of counters;
Last 3 counters have special installation

## Slide 42
Counters #46 & #47 are special assemblies that mount together at the inner cone

## Slide 43
Modified 48th counter installed

## Slide 44
Remaining Tasks
On Project
Off Project

## Slide 45: CTOF EVMS Chart
CTOF project 85.7% complete as of May 2015; 90% as of June 2015
>
≈

## Slide 46: CTOF Schedule
Act. ID | Activity Name | Start | Finish | Expected Finish | % Comp May
242231115 | Light Monitoring System Design | 2-Dec-13 | 31-Jul-15 | 31-Jul-15 | 75
242231260 | Final Drawings CTOF Support Harness | 3-Jan-14 | 31-Jul-15 | 15-Jul-15 | 70
242231265 | CTOF Procure Support Harness | 3-Aug-15 | 5-Oct-15 | 30-Sep-15 | 0
242231270 | Final Drawings Counter Installation Arm | 5-Jan-15 | 30-Sep-15 | 21-Aug-15 | 50
242231272 | Counter Installation Arm – Engineering Analysis | 4-May-15 | 31-Aug-15 | 31-Jul-15 | 33
242231275 | Procure Installation Tooling and Transport Rack | 1-Oct-15 | 3-Dec-15 | 11-Nov-15 | 0
242231300 | Light Monitoring System | 6-Oct-14 | 14-Aug-15 | 31-Aug-15 | 13
242231325 | Assemble Setup for Magnetic Shield Tests | 7-Apr-14 | 22-Jun-15 | 20-Jun-15 | 80
242231335 | Purchase Cables/Connectors (LMS Procure) | 5-Jan-15 | 31-Aug-15 | 15-Jul-15 | 53
242231340 | Fabricate Cables LV and HV | 6-Oct-14 | 30-Jun-15 | 15-Jul-15 | 50
242231355 | Correction Coil Power Harness | 6-Oct-14 | 30-Jun-15 | 30-Jun-15 | 33
242231360 | Fabricate Splitter Parts | 6-Oct-14 | 30-Jun-15 | 15-Jul-15 | 50
242231365 | Assemble CTOF Patch Panels | 15-Jun-15 | 30-Jun-15 | 31-Jul-15 | 0
242231370 | Assemble CTOF Splitters | 26-Jun-15 | 30-Jun-15 | 31-Jul-15 | 0
242231385 | Testing External Magnetic Shields | 13-Apr-15 | 17-Jul-15 | 20-Jun-15 | 32
242231415 | Assemble the DAQ System | 13-Apr-15 | 30-Jun-15 | 15-Jun-15 | 65
242231420 | Standalone CTOF Testing | 3-Aug-15 | 21-Aug-15 | 24-Aug-15 | 0
242231425 | CTOF Storage Rack for Transportation | 4-Dec-15 | 13-Jan-16 | 24-Nov-15 | 0
242231430 | Incidentals | 2-Jul-14 | 1-Jul-15 | 24-Nov-15 | 38
242231470M | CTOF Ready for Installation |  | 13-Jan-16 |  | 

## Slide 47: CTOF Schedule
Act. ID | Activity Name | BL Start | BL Finish | Expected Finish | % Comp May | % Comp June
242231115 | Light Monitoring System Design | 2-Dec-13 | 31-Jul-15 | 31-Jul-15 | 75 | 75
242231260 | Final Drawings CTOF Support Harness | 3-Jan-14 | 31-Jul-15 | 15-Jul-15 | 70 | 90
242231265 | CTOF Procure Support Harness | 3-Aug-15 | 5-Oct-15 | 30-Sep-15 | 0 | 0
242231270 | Final Drawings Counter Installation Arm | 5-Jan-15 | 30-Sep-15 | 21-Aug-15 | 50 | 65
242231272 | Counter Installation Arm – Engineering Analysis | 4-May-15 | 31-Aug-15 | 31-Jul-15 | 33 | 33
242231275 | Procure Installation Tooling and Transport Rack | 1-Oct-15 | 3-Dec-15 | 11-Nov-15 | 0 | 0
242231300 | Light Monitoring System | 6-Oct-14 | 14-Aug-15 | 31-Aug-15 | 13 | 13
242231325 | Assemble Setup for Magnetic Shield Tests | 7-Apr-14 | 22-Jun-15 | 20-Jun-15 | 80 | 100
242231335 | Purchase Cables/Connectors (LMS Procure) | 5-Jan-15 | 31-Aug-15 | 15-Jul-15 | 53 | 53
242231340 | Fabricate Cables LV and HV | 6-Oct-14 | 30-Jun-15 | 15-Jul-15 | 50 | 100
242231355 | Correction Coil Power Harness | 6-Oct-14 | 30-Jun-15 | 30-Jun-15 | 33 | 100
242231360 | Fabricate Splitter Parts | 6-Oct-14 | 30-Jun-15 | 15-Jul-15 | 50 | 100
242231365 | Assemble CTOF Patch Panels | 15-Jun-15 | 30-Jun-15 | 31-Jul-15 | 0 | 100
242231370 | Assemble CTOF Splitters | 26-Jun-15 | 30-Jun-15 | 31-Jul-15 | 0 | 100
242231385 | Testing External Magnetic Shields | 13-Apr-15 | 17-Jul-15 | 20-Jun-15 | 32 | 100
242231415 | Assemble the DAQ System | 13-Apr-15 | 30-Jun-15 | 15-Jun-15 | 65 | 100
242231420 | Standalone CTOF Testing | 3-Aug-15 | 21-Aug-15 | 24-Aug-15 | 0 | 50
242231425 | CTOF Storage Rack for Transportation | 4-Dec-15 | 13-Jan-16 | 24-Nov-15 | 0 | 0
242231430 | Incidentals | 2-Jul-14 | 1-Jul-15 | 24-Nov-15 | 38 | 38
242231470M | CTOF Ready for Installation |  | 13-Jan-16 |  |  | 

## Slide 48: CTOF Remaining Manpower - May

## Slide 49: CTOF Remaining Manpower - June
WBS 1.4.2.2.3.1 – 24.75 m-w remaining
WBS 1.4.2.6.4 - 11.2 m-w remaining

## Slide 50: Task #1 - CTOF Procurement I
Solenoid Support Structure:
  Activity ID(242231265); budgeted amount $46,930
  Estimated amount: $59,340
    Upstream support: $41,340 (RFI submission, Vision Machine)
    Support Ring + Hardware: $18,000 (engineering judgment)
  Conceptual design completed (Cyril Wiggins)
  Engineering analysis completed (Saptarshi Mandal)
  Detailed design nearly complete based on significant work to prepare RFI drawing package in May.
Procurement will send drawing package out for bids by the end of June (part #1) and mid-July (part #2).
Expect award of contracts by beginning of August and deliveries by mid-September

## Slide 51: Task #2 - CTOF Procurements II
Installation Tooling and Fixtures:
  Activity ID(242231275); budgeted amount $11,353
  Estimated amount: $11,353 (engineering judgment)
  Conceptual design and engineering analysis in progress (Bob Miller)
  Detailed design to be completed (Cyril Wiggins)
  Design & procurement to begin after submission of support structure

## Slide 52: Task #3 - CTOF Procurement III
Light Monitoring System:
  Activity ID(242231300)/(242231335); budgeted amount $27,176
  Estimated amount: $48,705
    Fiber bundle assembly: $39,705 (vendor quote for current version)
    LED driver/source: $2,000 (previous procurement HTCC)
    Integrating sphere: $3,000 (previous procurement HTCC)
    Fiber mounting blocks: $2,000 (engineering judgment)
    Hardware: $2,000 (engineering judgment)
    Conceptual design completed (Cyril Wiggins)
    Worked with local company on design and quote (LEONI Fiber Optics, Williamsburg)
    Submit PR for fiber mounting blocks by mid-July; remainder of procurements
    by mid-August
LED driver

## Slide 53: Task #4 - CTOF Procurement IV
Modify Counter #48:
  Activity ID(242231430); budgeted amount $2,000
  Estimated amount: $2,000 (engineering judgment)
  Machining work to be completed by JLab or outside Machine Shop
  Machine wedge-shaped counter into rectangular-shaped counter to allow for insertion in the barrel
- Machine shop work to begin by end of July
Counter then needs to be glued, polished, and wrapped
Counter shields and PMTs to be installed
Counter then to be inserted into test stand
(DSC, Gegham Asryan)

## Slide 54: Task #5 - CTOF Procurement V
Transportation Crates to Hall B:
  Activity ID(242231425); budgeted amount $17,719
  Estimated amount: $17,719 (previous procurement; engineering judgment)
Design drawings completed
Still need to finalize detailed plans for transport of counters to Hall B
Preparations for procurement are last on the CTOF work list given priorities.

## Slide 55: Task #6 - CTOF Cosmic Ray Studies
Cosmic Ray Testing:
  Activity ID(242231420); 1.5 m-w remaining
  Work Plans:
- Study system noise, baseline fluctuations, and stability
HV gain matching of all PMTs using mean ADC spectra
Measurement of average counter time resolution using counter triplets
Investigation of and attention to problematic system elements
Optimization of signal splitter ratios (presently 90/10)
Status:
Claimed as 50% done for June progress
Initial system checkout and gain matching complete
Working to measure initial counter time resolutions
Will complete work by end of July
(DSC, Andrey Ni)

## Slide 56
Remaining Tasks
On Project
Off Project

## Slide 57: CTOF Off-Project Proposed Work
The cosmic ray testing plans that will take place off project include:
Detailed studies of time resolution vs. coordinate and angle for all counters
Studies to minimize CFD contribution to time resolution
Time-dependent calibration optimization of time resolution
CTOF calibration suite development (in progress)
CTOF reconstruction suite development (in progress)
Preparation of technical reports/papers
Development and optimization of light monitoring system
Long base-line monitoring of system performance and stability
Calibration development off-project similar in scale to work on FTOF system
20% FTE scientist level
(DSC)

## Slide 58: CTOF Project Status Summary
The CTOF detector assembly is very nearly complete.
    - Cosmic ray studies are now in progress
The design and engineering work related to CTOF is now coming to its end.
- Solenoid support structure - Light monitoring system
- Installation fixtures and tooling - Counter #48 work
Final procurements are in progress and costs are now well understood. All PRs should be submitted by the end of August.
Project budget: $106,178 Current estimates: $140,117 (+32%)
Scheduled for close-out of WBS 1.4.2.2.3.1 by January 2016; On-project assembly and testing to be completed by end of August 2015.
Cosmic ray testing will continue until the end of calendar year 2015
(off project – 20% FTE) for fine tuning of calibration parameters and understanding LMS settings/operation.
Detector installation scheduled for summer/fall 2016.