## Slide 1: A Test Cavity & Cryostat for Rapid RF Characterization of Superconducting Materials
Paul B. Welander, Matt Franzi, Sami Tantawi
SLAC National Accelerator Laboratory, Menlo Park, CA 94025
28 July 2016

## Slide 2: Superconducting RF Materials
Collaborate w/ others
in DOE complex and
beyond to advance
state of the art for
SRF materials.
Example: Temple MgB2
  RF measurement as tool
  for process development
Recent measurements:
  MgB2 – Temple Univ., Peking Univ.
  Nitrides – Naval Res. Lab, MIT Lincoln Lab
  Nb – Alameda Applied Sciences, JLab
SRF Materials Development
Two hemispherical test cavities (one Nb, one Cu) to measure surface resistance & quenching field at 4 K.
X-band operation (11.4 GHz)  small sample size, 33% of cavity loss from sample surface.
Closed-cycle, pulse-tube cryocooler enables 24-hr. test cycle  unmatched throughput & rapid feedback.
2” dia. sample
Novel SRF Cavity Fabrication
Develop SRF coatings that can be applied to novel accelerator structures being designed and built at SLAC.
Optimize coatings & cavity design for:
  High Efficiency – higher Q, lower dynamic loss, less cooling power required
  High Gradient – higher beam energy
  High Temperature – operation at 4 K
Optimize Q0

## Slide 3: System Capabilities
  SLAC test cavities and cryostat enable rapid (24-hr. cycle) characterization of superconducting RF (SRF) materials.
  Characterize surface impedance by measuring the quality factor, Q0, of a cavity at 11.424 GHz, down to 4 K.
Capable of low power (PNA) and high power (Klystron) measurements.
Compact design thanks to X-band operation (5.6” diameter).
Interchangeable flat cavity bottom, fits 2” (50.8 mm) diameter samples up to 0.25” (6.25 mm) thick.
Cavity design maximizes H-field and minimizes E-field on the sample surface.
Cu and Nb cavities allow us to measure surface resistance (Rs), quenching field (Hquench), and transition temperature (Tc).
Can achieve Hpeak ~ 360 mT with 50 MW Klystron.

## Slide 4: Cryomech Pulse-Tube Cryocooler
Our cavity cryostat utilizes a Cryomech cryorefrigerator.
Two-stage pulse-tube operation
Base temperature of 3.5 K with cooling power of 1.35 W at 4.2 K
Utilize the remote motor version to minimize cavity vibrations.
First stage (40 K) used for thermal shielding and cold section of waveguide.

## Slide 5: Cavity Cryostat Assembly – Model View
Cryocooler
2nd Stage
Sample
Under
Test
RF Feed
40 K Shield
Diode
Temp
Sensors
Sample
Plate
Cavity Iris

## Slide 6: Hemispherical Cavity Design – HFSS Modeling
High-Q hemispheric cavity with a TE032-like mode at 11.4 GHz
  Maximum H-field (2.5x Hdome), zero E-field on sample
  Sample accounts for 8% of cavity area, but 33% of cavity loss
  No radial current on the cavity bottom
r = 0.95”
RF Feed

## Slide 7: Nb-Coated Cavity Design
Sample Surface
Hemisphere Surface
f0 = 11.4 GHz
Qtotal = 1.6e7
Gtotal = 1416 Ω
GNb = 2120 Ω
Gsample = 4264 Ω

## Slide 8: Two Cavities
Coated w/ 5 μm Nb film at CERN (S. Calatroni)

## Slide 9: Cavity Assembly
Sample
Under
Test
Sample
Plate
Cavity Iris
RF Feed
40 K Shield

## Slide 10: System Photo and RF Measurement Network
Cryostat
Waveguide to Klystron/NWA
Measurement ports:
Forward Power: 5 (and 2)
Reflected power: 4
(Waveform measured by either a peak power meter or a scope with mixers)
Low-power PNA measurement: 3 (or 6)
1
2
3
4
Cavity
Klystron
10dB
45dB
45dB
5
6
55dB
7
Cryostat
Mode converter
Bend
Load
System Diagram

## Slide 11: Bulk Nb Reference Sample
Single-crystal bulk Nb from DESY
  Received January 2008
  Baked in 2010, untreated since
Q0 in Cu limited by cavity materials
In Nb cavity at 4 K, Q0 translates to Rs = 65 μΩ
  Assumes Rs,sample = Rs,cavity
  Standard deviation of 1%
  Assuming f 2 and (T/Tc)4 dependence,  Rs = 47 nΩ at 2.0 K and 1.3 GHz
Q0 vs T for Nb Reference
in both Nb & Cu cavities

## Slide 12: Nb Films from AASC & JLab
Low power measurements in our Nb cavity.
Nb films on copper (JLab, A.-M. Valente-Feliciano) & stainless steel (AASC, K. Velas) compare favorably with our bulk Nb sample.
Assuming a cavity Rs of 65 μΩ, both films have Rs of about 17 μΩ.

## Slide 13: MgB2 on Copper from Temple Univ.
Series of MgB2 films grown on copper last summer at Temple Univ. (W. Withanage, X. Xi).
Q0 measurements served as feed-back to develop growth process, enabling rapid improvement.
Tc’s up to 38 K were measured in Cu cavity.

## Slide 14: MgB2 on Niobium from Peking Univ.
Recently measured two MgB2 films grown on niobium at Peking Univ. (Z. Ni, K. Liu).
Process improvement over past eight months, reducing Rs ~ 1 OM

## Slide 15: Cavity Cryostat Status & Summary
Cu and Nb cavities allow us to measure surface resistance (Rs), quenching field (Hquench), and transition temperature (Tc).
Low-power Q vs. T takes less than 24 hrs.  rapid feedback for film growth process development.
Currently building up capability to perform high-power testing, and measure Hquench.
Built in concurrent capability to measure samples at low power in both cavities.

## Slide 16: Highly Efficient Direct-Feed Split-Cell Cavity
Z. Li & S. Tantawi

## Slide 17: Demonstration in Cu at X-band
20-cell X-band structure fed by two waveguides from a single RF input.
Bead-pull measurement shows uniform field dist.
Currently under test, has exhibited up to 130 MV/m.

## Slide 18: Adapting for SRF
Direct-feed cavity utilizes highly reentrant cell shape, shifting max-H from equator.
RF cavity loss reduced by nearly 60% c.f. TESLA.
 | 1.3 GHz TESLA | 1.3 GHz direct feed
R/Q (ohm/m) | 984.0 | 2571.4
Esurf/Eacc | 2.02 | 5.32
Bsurf/Eacc (mT/(MV/m)) | 4.17 | 4.04
Ploss (W/m/(MV/m) | 0.101 | 0.043
Q0 | 1e10 | 0.91e10

## Slide 19: Tuning Isolated Cells
Tuning Isolated Cells
Plunger w/ zero offset from the coaxial center

## Slide 20: An Efficient SRF Split-Cell Cavity
Challenge # 1 is how to fabricate:
  Complicated structure precludes PVD – only vapor-phase dep seems plausible.
  Cu structures are brazed. SRF cavity to be welded or bolted.
  Accelerating mode has no azimuthal current, but excitation of HOMs and a lossy joint could kill efficiency.
Current plan is to fabricate a 2-cell S-band structure from bulk Nb:
  Measure low-power Q
  Demonstrate tuning