## Slide 1: Progress of the Commissioning of the Choked Resonator Cavity for SRF Thin Film Measurements
Philippe Goudket
on behalf of the collaboration team
ASTeC RF Group,
STFC Daresbury Laboratory, UK

## Slide 2: Motivation
There is an considerable demand for RF Thin Film testing capacity.
We require a system allowing a reasonable turnover of samples in order to serve our TF deposition program.
Initial plans call for the deposition and testing of planar samples.

## Slide 3: Design criteria
Avoidance of joins, seals and gaskets
  To increase repeatability and reduce preparation time
Separate cavity and sample
  To allow thermometry and to allow for quick sample rotation
Sample to be as simple as possible
  To minimise cost and complexity. All sample support to be done by a dedicated plate
Initial frequency choice: 3.9 GHz
  As it is a frequency used by real SRF cavities while also restricting dimensions.

## Slide 4
Original idea
Top Plate (Al)
Bottom Plate (Al)
180.0
29.23
38.0
19.23
19.23
25.0
4.0
8.0
50.0
4.0
20.23
3.60
Thru hole
50.63
50.63
69.86
3.9 GHz cavity

## Slide 5: Refinement of the design
3.9 GHz cavity
Multipactor optimisation
7.8 GHz cavity
Sample size limit in deposition chamber

## Slide 6: RF optimisation of the design
Optimisation steps:
Sharp S11on resonance
Minimal S21 while showing a narrow bandwidth resonance peak
The lowest possible edge field
The lowest possible absorbed power in the absorber/ loss minimisation

## Slide 7: RF optimisation of the design
Optimisation steps:
Low S11on resonance
Minimal S21 while showing a narrow bandwidth resonance peak
The lowest possible edge field
The lowest possible absorbed power in the absorber/ loss minimisation

## Slide 8: RF optimisation of the design
Optimisation steps:
Low S11on resonance
Minimal S21 while showing a narrow bandwidth resonance peak
The lowest possible edge field
The lowest possible absorbed power in the absorber/ loss minimisation

## Slide 9: RF optimisation of the design
Optimisation steps:
Low S11on resonance
Minimal S21 while showing a narrow bandwidth resonance peak
The lowest possible edge field
The lowest possible absorbed power in the absorber/ loss minimisation

## Slide 10: Measurement method (RF)
G, ps and pc determined by simulation

## Slide 11: Measurement method (Calorimetry)
More accurate method
- Can reach sub nW resolution
Courtesy B.P. Xiao
Sample temperature
Thermometry

## Slide 12: LLRF Control System

## Slide 13: Cavity reoptimisation - 3rd choke added
Normalised electric field intensity measured through gap between cavity and sample
Normalised E in gap
E-field

## Slide 14: Field Distribution
E-field
H-field
37%
63%

## Slide 15: Two and three choke cavities
Fine grain
RRR 200 Nb
Large grain
RRR 400 Nb
Cavities manufactured by

## Slide 16: Cryogenic and support design
Liquid Helium
Chamber
Liquid Nitrogen
Chamber
Outer Vacuum
Chamber
77K Radiation Shield
Sample Chamber
Cradle Assembly
RF cables and thermometry wiring
Schematic of the cryostat.
Solid copper post (cavity support)
Hollow stainless steel post (sample support)
Tuneable Input Coupler

## Slide 17: Cradle assembly
UP
DOWN

## Slide 18
>> Assembly Count down to experiments
How it came together in 10 months (in house)
Oct 2014 >> Design
July 2015
Manufacturing >>
>> Quality Control >>

## Slide 19: And then…
First cryogenic test
  Vacuum leak – temperature not reached
Second cryogenic test (LN2)
  Verification of the vacuum seal – OK!
  Issues with the transmission probe
    Attempt to use a patch antenna – space insufficient in gap
    Modification of the side-coupling probe
  Issues with the tuneable input coupler
    Design adapted

## Slide 20: And now…
Reassembly
  RF leakage noticed in inner vacuum chamber
    Room temperature leakage estimated to have a Qe ~ 22000
    Identification of leakage mechanism still under investigation!
  Radiation shielded still needs commissioning
On the horizon
  Measure cavity performance with cold test
    Proceed with scientific programme
    (or return to the drawing board)

## Slide 21: The UK’s SRF collaboration team
S. Wilde, Dr. B. Chesca
P. Pizzol, Prof. P. Chalker
F. Lockwood-Estrin, Prof. J. Bradley
L. Gurran, Dr. G. Burt
Dr. R. Valizadeh, Dr. O.B. Malyshev,
Dr. P. Goudket, S. Pattalwar,
A. Hannah, N. Pattalwar, K. Dumbell
J. Herbert, A. Wheelhouse,
P. McIntosh
S. Bibby-Trevor, P. Davies,
E. Jordan, T. Jones
Dr. G. Stenning