## Slide 1: Update on Silicon Photomultipliers
Yi Qiang
DOE S&T Review
May 10, 2012
Notes: Good morning
In this talk, I’m going to give you some update on our recent study on the photomultipliers, particular related with its applications in Hall-D.

## Slide 2: Bio – Yi Qiang
MIT – Ph.D. (2007)
  Search for pentaquarks in Hall-A at Jefferson Lab
Duke University – Post. Doc. (2007 – 2010)
  6 GeV neutron transversity in Hall-A at Jefferson Lab, in charge of polarized 3He target
  Co-spokesperson of a 12 GeV SoLID SIDIS experiment using a longitudinal polarized 3He target
Jefferson Lab Hall-D – Staff Scientist (2010 – present)
  In charge of photon beam radiator
  Silicon photo multipliers, particularly for the Barrel Calorimeter
Notes: My name is Yi Qiang.
Both of my Ph.D. study at MIT and post doc work at duke university are mainly related with experiential nuclear physics research in JLab Hall A, including the search for pentaquarks and the first measurement of neutron transversity using a polarized 3He target.
In 2010, I become a staff scientist of Hall-D. And my main responsibilities are the photon beam radiator and the silicon photo multipliers.

## Slide 3: Introduction
Silicon Photo Multiplier (SiPM) is a new type of photon-counting device made up of multiple Avalanche Photo-Diode (APD) pixels operating in Geiger mode. Each APD pixel outputs a pulse signal when it detects one or more photons, and the output of the SiPM is the total sum of all these pulses.
Hamamatsu
4×4 Array of
3×3 mm2 SiPMs
50 mm
12 mm
Notes: Made of many avalanche photo diode pixels.
Each APD outputs a pulse signal when a photo electron is created and induces avalanche.
As all the pixels are working in Geiger mode, the summed output of a SiPM detector is proportional to the number of pixels fired.
Here is the a photo of a SiPM array we are going to use in Hall-D as the photon detector of the barrel calorimeter. Each array consists of 16 3x3 mm2 SiPMs. With a maginfying glass, you will be able see its 50 um pixels.

## Slide 4: Features of SiPM
Immune to strong magnetic field: few Tesla
Good photon detection efficiency (PDE) : > 20%
High gain: ~ ×106
Compact size
No HV needed
Photon counting for weak light
Some disadvantages
  Noise rate
    can be reduced by cooling the sensor
  Cross talk and after pulses: 10 ~ 20%
  Limited range of gain:
    strong corr. between gain and PDE
  Radiation hardness
1 pixel
2 pixels
2 pixels
1 pixel
Amplitude
Time
20 ns
Notes: Arrows, APD few hundreds gain, pulse shape
List features
Top plot: typical signal output captured by an oscilloscope, bottom plot, integrated charge distribution. You can clearly identify the peaks corresponding to different numbers of pixels fired.
Radiation hardness is the major concern to use SiPMs in nuclear and high energy physics. Therefore, at Jefferson lab, we did a lot of study of the radiation damage and I will show you more details later in my talk.

## Slide 5: Time Resolution and Recovery Time
PLP-10 Picosecond Light Pulser from Hamamatsu:
  Pulse width < 70 ps
  Repetition rate up to 100 MHz
Intrinsic time resolution
  < 30 ps
  Good enough for most of
  the TOF applications
Pixel recovery/dead time
  10 ~ 50 ns, depends on pixel size,
  smaller size → shorter recovery
  Can handle MHz level input rate
A NIM article is being prepared
25 mm pixel
50 mm pixel
Resolution
of TDC
low → high
input light intensity
50 mm pixel
Amplitude of the 2nd pulse
Time difference between two pulses (ns)
50 ns dead time
Notes: Other than the features I previously mentioned, in Hall-D we also studied the timing properties of SiPMs to find its feasibilities in other applications.
In order to do this study, we purchased

## Slide 6: SiPM for Hall-D Barrel Calorimeter
Barrel Calorimeter (BCAL)
  Scintillator fiber – Lead
  48 trapezoidal modules
  Photon energy and position
Reasons to use SiPM
  Strong magnetic field: > 1 T
  Tight space for readout
  Similar cost as fine-mesh PMTs
  Good time resolution
Additional benefits
  Natural radial segmentations for additional PID
BCAL
Barrel Calorimeter
Photon beam
CDC
Central Drift Chamber
FDC
Forward Drift Chambers
Future
Particle ID
TOF
Time-of-Flight
Solenoid: 2.2 Tesla
FCAL
Forward
Calorimeter
BCAL
Module
Light Guides
SiPMs
phone on the other end
Sci. Fiber
Lead
Notes: Such a configuration gives us additional PID capability for the natural radial segmentations.

## Slide 7: Applications for Scintillating Counters in Hall-D
Beam line hodoscopes: tagger microscope and pair spectrometer
  Small size to couple to thin fibers
  Capability to handle high
  rates, short recovery time
Start Counter:
  Small size
  Good time resolution
SiPMs and Electronics
Scintillating Fibers
Photon Tagger Microscope
SiPM
Bent Scintillator
Source and
Trigger PMT
Start Counter Test Setup
Notes: Since the SiPMs have the great capablity to deal with high couting rate, we are going to use them in our photon beam hodoscopes, including the tagger microscope and the pair spectrometer.
In addtion, we will also use the sipms for the start couter as part of our time of flight system for its extraordinary time resolution.

## Slide 8: Radiation Damage to Silicon Detectors
Bulk radiation damage in silicon detectors is primarily due to displacing a Primary Knock on Atom out of its lattice site.
Proportional to The Non-Ionizing Energy Loss (NIEL).
Heavier particles induce greater damage.
Neutron damage is the major concern for SiPMs in Hall D:
  > 90% damage
  10 years running: 2.5×109neq/cm2
  Requires dark rate < 100 MHz for each SiPM array

## Slide 9: Neutron Radiation Tests at JLab
Test using neutron background during P-REX in Hall-A
  12×12 mm2 SiPM arrays from Hamamatsu and SenSL, 1×1 mm2 and 3×3 mm2 SiPMs from Hamamatsu (50 mm)
  Self-annealing at room temperature
  Need to cool these devices during GlueX operation
Test using JLab RadCon AmBe source
  1×1 mm2 SiPMs from Hamamatsu (50 mm)
  Measure temperature dependence
    Radiation damage
    Self-annealing
    Impact of radiation damage on the dark rate temperature dependence
  Accumulated radiation damage with annealing
NIM paper being submitted

## Slide 10: First Look at the Radiation Damage
10% drop in amplitude
No change in width
Damage independent
on bias voltage
Approximate Linear damage curve
2 days during P-REX@Hall-A
15 years
@Hall-D

## Slide 11: Annealing at Room Temperature
About half of the increase of the dark current caused by radiation damage recovered
Decay time ~ 10 days
Drop in amplitude fully recovered

## Slide 12: Test Sequence using AmBe Source
12 1×1mm2 Hamamatsu SiPM (50 mm)
Different temperature conditions/combinations
Two irradiations and two annealings
Total fluence ~ 13 years of high intensity running of GlueX
1st Irradiation
1st Annealing
2nd Irradiation
2nd Annealing
25oC ×6
-5oC ×6
60oC ×4
25oC ×4
-5oC ×4
60oC
×8
25oC ×2
-5oC ×2
25oC ×6
-5oC ×6
40oC ×12
Numbers of units tested at different temperatures are shown in the blocks

## Slide 13: Recovery at Different Temperatures
Higher Temperature brings faster and stronger self-annealing
Recovery baseline:
saturates at higher temperature
Fit curve: exponential decay
Results from first irradiation
Exponential dependent
decay time

## Slide 14: Damage Curve
Radiation damage does NOT depend on temperature
Radiation damage does NOT depend on previous irradiations
1×1 mm2 SiPM@25oC w/ annealing
1st irradiation
2nd irradiation
Equivalent to 13 years of operation

## Slide 15: Expected Lifetime in Hall-D
Chart: High Intensity Operation in Hall-D (year)
| 0.0 | 3.696
Prototype | 14.5778 | 59.3778
Production Unit 01 | 4.0533 | 64
Production Unit 03 | 3.8756 | 58.6667
Hamamatsu improved initial dark rate of the production units, still same damage rate as prototype samples.
Life time in Hall-D ~ 8 years with high intensity beam:
  108 g/s on a 30 cm LH2 target, 30% running efficiency
  SiPMs cooled to 5oC, periodic annealing at 40oC
8 years@5oC
with periodic
annealing at 40oC
12×12 mm2 SiPM Array
BCal Limit
Notes: Actually, after Hamamatsu sent us their prototype a couple of years ago, they managed to reduce its dark rate when they deliver our BCAL production units. Although such a change does not effect the damage slope, it indeed brings more margin for radiation damage.

## Slide 16: R&D for Electron-Ion Collider
Except for the radiation hardness, SiPMs are ideal for future high energy detector systems, such as EIC:
  High gain, compact form factor, good photon detection efficiency and immunity to magnetic field.
An R&D proposal “to test improved radiation tolerant Silicon Photomultipliers” for EIC has been approved:
  Collaborate with JLab detector group and Hamamatsu. Perform bench tests on samples with varied manufacturing parameters, then give feedback to Hamamatsu and continue the loop.
  Received the first group of samples in April, test is ongoing.

## Slide 17: Summary and Outlook
Silicon photo multiplier is a novel photo detector with many desirable features.
We thoroughly studied its performance including radiation hardness, time resolution and dead time.
Applications in Hall-D:
  Barrel Calorimeter
  Beam line hodoscopes
  Start counter
Future applications in EIC
  Many possibilities
  R&D with Hamamatsu on the radiation hardness

## Slide 18: Acknowledgement
Hall-D
  Elton Smith (Scientist)
  Alexander Somov (Scientist, SiPM timing measurement)
  Ivan Tolstukhin (Grad. Student, SiPM timing measurement)
Detector Group
  Carl Zorn (Scientist, SiPM characterization and radiation measurements)
Radiation Control Group
  Pavel Degtiarenko (Scientist, Flux simulation)
  Melvin Washington (Radiation monitoring)
  John Jefferson (AmBe source)