## Slide 1: Energy (ILC) and Intensity (Project X) SRF Cavity Needs
C.M. Ginsburg (FNAL)
SRF Materials Workshop
Jefferson Lab
July 16-17, 2012

## Slide 2: International Linear Collider
e- e+
ILC Specifications
Initial center-of-mass energy up to 500 GeV, upgradeable to 1 TeV
Ability to scan in energy between 250 and 500 GeV
Integrated luminosity in the first four years 500 fb-1 at 500 GeV
  Peak luminosity 2 x 1034 cm-2 s-1, 75% duty factor
Main Linac
Two 11-km linacs, one for electrons and one for positrons
Accelerate from injected energy of 15 GeV to the final beam energy of 250 GeV
~16000 1.3 GHz superconducting RF cavities, operating at average gradient 31.5 MV/m
ILC RDR
Notes: This design information is from the ILC Reference Design Report which was released last month.

## Slide 3: ILC: SRF Cavities
Primary design criteria
  Maximum gradient in shortest distance
  Minimize cost per cavity (or per cryomodule)
TESLA shape was chosen as baseline
  Epeak/Eacc = 2.0
  Bpeak/Eacc = 4.26 mT/(MV/m)
Cavity performance specification
  vertical (qualification) test: 35 MV/m and Q0>0.8E10
  operation: 31.5 MV/m and Q0>1E10
Required cavity processing
  Electropolishing, HPR, class-10 cleanroom assembly

## Slide 4: ILC: Production Yields
1st-pass cavity yield at >35 MV/m is (31 +- 7) %
2nd-pass cavity yield at >35 MV/m is (57 +- 8)%
KILC 2012
1st pass
KILC 2012
2nd pass
Electropolished 9-cell ILC cavities built by established vendors
Processed by established labs: JLab/DESY/KEK
Max gradient [MV/m]
Max gradient [MV/m]
Yield [%]
Yield [%]
100
100
0
0
>25
>30
>35
>25
>30
>35

## Slide 5: ILC: Beyond the Baseline Studies
50 MV/m
  Several instances of this gradient level with different cell shapes
    trade off lower Bpeak/Eacc = ~3.5 mT/(MV/m) for higher Epeak/Eacc ~2.2
  Avoiding field emission (even) more important
Reduce cost
  Hydroformed cavities? (shapability, Nb properties)
  Large grain material? (surface roughness, earing, BCP only)
  Looser material specs? (tantalum content, targeted inspection)
  Better performance after first pass (Tumbling? Better diagnostic tools?)

## Slide 6: Project X at Fermilab
High-power proton beam with flexible beam formats delivered simultaneously to multiple users
  3 MW at 3 GeV
    Rare process experiments with kaons, muons, nuclei
    Extraction of up to 0.5 mA at 1 GeV is possible
  0.080-0.200 MW at 8 GeV
    Rare process experiments, neutrinos
  2 MW at 60-120 GeV
    Long-baseline neutrino experiments
Path toward neutrino factory or muon collider

## Slide 7: Project X Linac SRF Cavities
b=0.11
b=0.22
b=0.4
b=0.61
b=0.9
325 MHz
10-160 MeV
b=1.0
1.3 GHz
3-8 GeV
ILC
162.5 MHz
2.1-10 MeV
650 MHz
0.16-3 GeV
Section | Freq | Energy (MeV) | Cav/mag/CM | Type
HWR (G=0.11) | 162.5 | 2.1-10 | 9 /6/1 | Half Wave, solenoid
SSR1 (G=0.22) | 325 | 10-42 | 16/8/ 2 | Single Spoke, solenoid
SSR2 (G=0.47) | 325 | 42-160 | 36/20/4 | Single Spoke, solenoid
LB 650 (G=0.61) | 650 | 160-460 | 42 /14/7 | 5-cell elliptical, doublet
HB 650 (G=0.9) | 650 | 460-3000 | 152/19/19 | 5-cell elliptical, doublet
ILC 1.3 (G=1.0) | 1300 | 3000-8000 | 224 /28 /28 | 9-cell elliptical, quad
CW
Pulsed
H-
Notes: 6 cavities/CM for low-beta 650’s
8 cavities/CM for high-beta 650’s

## Slide 8: Project X: SRF Cavities
Primary design criteria
  Minimize power dissipation (esp. 650 MHz b=0.9)
Performance challenges
  Requirements for Epeak and Bpeak are modest
    Multipacting is an issue for the complex shapes: cleanliness
  Power dissipation requirement on 650 MHz cavities must be optimized
    JLab design for b=0.61 achieved proof-of-principle Q0 in single-cells using BCP
    No data for b=0.9 yet, but Q0 ~2E10 in the 15-20 MV/m range, in operation, may be challenging
      How to optimize Q0

## Slide 9: General Materials Issues
For any of our projects…

## Slide 10: Features seen on an ILC cavity
before bulk EP
Cavity was light BCP’d (10-15 um) at AES
after bulk EP at ANL/FNAL
1mm
TB9AES013

## Slide 11: More…
TB9AES013: Pits observed in all three images, but generally enhanced by EP. Pits are not restricted to just the equator weld or the heat affected zone.
1) Optical inspection of equator weld before EP
2) Photo before electropolishing.
3) Photo after electropolishing ( ~ 120 microns removed)
1
2
3
This cavity reached 38 MV/m!

## Slide 12: Niobium for Project X Cavities
An SSR1 made from this material performed the best ever!
Inspection by Z.Conway, T. Reid (ANL)
3mm Nb sheets
All surfaces are rough, were high-speed buffed
Lots of inclusions, pits
Indications of poor handling, e.g., Nb on Nb rub marks
Deep scratches

## Slide 13: General Materials Issues
What types of material flaws matter?
  Material cost, inspection/remediation effort, risk, etc.