## Slide 1: LERF Readiness Review
Kevin Jordan
LERF Cryomodule Test Facility Overview
November 12 - 13, 2018

## Slide 2: Outline
  Scope of the LCLS-II conversion
  Work planning
  Top level schedule
  Safety systems & radiation shielding
  Software, controls & network infrastructure
  Solid State Amps, LLRF & all that waveguide
  Cryomodule & tunnel connections
  Cryogenic connections
  Summary

## Slide 3: Scope of the LCLS-II Conversion Project
  Modify LERF to enable LCLS-II CM testing
  Take advantage of Safety System interlocks / shielded vault
  Install computer services; firewalled network & servers
    EPICS Software development & test bed for SLAC team
    PCs with PLC licenses to remotely develop cryo & vacuum PLC controls
  Install16 SSAs & reroute existing waveguide to LCLS-II modules
    1400 feet of new waveguide
  Install, test & commission LCLS-II production racks (SLAC staff)
    LLRF, Cryo Controls, SC Magnets (w/Quench protection) & Vacuum
  Enable testing and running of 2 CMs together, this simulates Linac 1
  Operate 2 cryo modules with a similar cryo system rather than CTF
    Gain invaluable experience cooling down 2 CMs at SLAC tunnel grade

## Slide 4: Work Planning
  Weekly schedule meetings are held in LERF breakroom
    Work planning for the upcoming week discussed
    Interferences discussed
  Long range schedule is posted on white board
    http://lerfboard.acc.jlab.org/board1.html
    http://lerfboard.acc.jlab.org/board2.html
  ATLis is used as work control document
    https://wiki.jlab.org/lerf/index.php/ATLis
  SAD is also used
    Controlled by Steve Suhring
  SOP for testing, travelers, & procedures
    Controlled by Mike Drury

## Slide 5: Top Level Schedule
The test schedule is drive by various scheduled shut downs of the physics program
Caveats
  CHL requires 200 watts + 10%; Either the LERF ¼ CM & FL02 are cold & full and/or the 2 LCLS-II CMs are cold and full
  The CHL 1 has to be at 4K for LCLS-II U-Tube operations
  The CHL 1 will be taken to 4 K in Dec. to manage Helium inventory for the first cool down in January. By going to 4 K early there is less of a time constraint for U-Tube operations
  First cool down is either complete or aborted by January 15 to enable hot check out of CEBAF prior to the February physics run
The test cycle is nominally 2 months
1St Break – January 2 to 15, that is followed by 2 weeks of hot check out & start of Physics Feb. 1
2nd Break – April 15 to May 10
3rd Break – August 15 to Sept. 15

## Slide 6: Safety System & Shielding
  Shielding assessment done for the vault with 2 LCLS-II modules, all 16 cavities at full gradient
    All 26 blocks will be in place for any testing
    Originally I hoped not to install & remove all 26 shield blocks each time but I gave up on that – it is not a cost driver with the crane already there
    Rad Group Doc: RCD-DEP-18#002 dated Jan 2018
    https://wiki.jlab.org/lerf/images/2/2b/LERF_TRUCK_ENTRANCE_SHIELDING.pdf
  PSS interlocks are identical to the CMTF in the test lab
    System duplicated from existing drawing to avoid new design effort
    Waveguides are pressurized (few psi) for standard interlocks
    PSS has already been certified, except for waveguide pressure

## Slide 7: Software & Controls
  Servers & network are up & on-line – Wesley Moore’s talk
  Two PCs are online with Siemens PLC licences
  Cryo-Con RTD controllers are up and running
    These are for the Cryo-can & non LCLS-II instrumentation
    New flow meters will be tied into cryo control system
    JT electric valve controls are re-assigned channels & cables are relocated to new locations (RT03, SH04, JT03)
  New high level aps for testing will be developed by JLab
    This is an opportunity for SLAC software development
    This will greatly speed the commissioning @SLAC
  Initial cooldown will be manually – Matt Marchlik’s talk
  Cavity Testing – Mike Drury’s talk

## Slide 8: SSAs & Gallery Waveguide Layout
  Entire system exists as a 3-D model in Solid Works
  Total project requires new waveguide ~380 pieces, ~ 1,400 feet
    + LCW cooling of circulator loads & SSAs
  Production LCLS-II SSAs, couplers, circulators, & loads are used
  All waveguide pieces fit (first time!)

## Slide 9: Solid State Amps
  ALL 16 SSAs have been run to full power
  All waveguide couplers have been calibrated
  Personnel Safety System (PSS) connections are complete

## Slide 10: Raised Floor in Gallery & Cable Tray
  The RF cable routing will be through the penetrations and along the gallery floor, protected by the raised floor
  This image is before the heliax cables have been installed, some sections of floor grating needed to be removed
  All cable & connectors are delivered
  Cryo rack cable installation ~99% complete (88 cables), all of those have been terminated in racks
  Tunnel terminations > 90 % complete

## Slide 11: LLRF Controls
Fully functional LCLSII LLRF control racks (4) are powered up & remotely accessible from SLAC
  Precision receiver chassis
  RF Station
  Resonance/Interlock chassis
  LO Chassis
  Power supply
  EPICS Interface
LLRF Rack

## Slide 12: Master Reference & LO Distribution in Progress
  1300 MHz Reference: Rohde & Schwarz Signal Generator
  1320 MHz LO: Derived from 10 MHz and 1300 MHz

## Slide 13: Tunnel Waveguide Layout
  Installation is complete in the gallery & tunnel

## Slide 14: Tunnel Waveguide Layout
  Installation complete

## Slide 15: Cryogenic Connections to Transferline
Cryo can was late 3 months, it was delivered on Nov. 1, placed Nov. 5, and now in process of having field fit joints tied in & instrumentation checked out
See Matt’s talk

## Slide 16: Process & Instrumentation Diagram
P&ID JL0043630
In approval process
Zone 00
Zone 01
Zone 02
Zone 03
CM 1
CM 2

## Slide 17: LCLS-II Cryogenic Controls & Interfaces
  JLab cryo staff controls cryo-can & JT valves for cooldown
  SLAC has access to CM JT valves & heater controls
  Two fully assembled sets of cryo-control racks are operational
  Tunnel terminations done by Jlab staff
    Rack ‘field’ terminations have been done by SLAC
  PLCs for cryo control are on line and can be remotely programed from SLAC
  Cryo racks will provide UPS power to the LLRF racks for backing out tuners during power failure
  SLAC staff encouraged to be present for cooldown

## Slide 18: LCLS-II Cryogenic Controls Racks (1 of 2 sets)

## Slide 19: Vacuum Controls & Interfaces
  Controllers for ion pumps and gauging from SLAC are on-site
    Gauges are part of CM assembly
  A LCLS-II vacuum pump cart will be shipped soon as a test of microphonics
    Roughing pump, turbo, pirani, cold cathode, & PLC controller
  Jlab provided pump cart will be initially used if SLAC cart is late
  EPICS control interface to pump cart, gauges & PLC
    Interlocks to LLRF

## Slide 20: Other Instrumentation
  Faraday cup on both ends
  Decarad monitors
  Fermi supplied Total Loss Monitor (TLM)
  Other radiation detectors
    Fox - ?
    Chipmunk - ?
    Scarecrow - ?
    TLDs - ?

## Slide 21: ES&H issues
  All radiation shielding blocks will be used for first test
    Dosimetry will be used between first & second layer to get a total dose
  Operational Safety Procedure is being finalized
    Mini review held for running SSAs into shorts
  ODH Assessment done, needs to be signed
    Less LHe volume than used in the past
  Pressure system design authority is Matt Marchlik
  Egress has been reviewed by Facilities
    https://wiki.jlab.org/lerf/index.php/Facilities_Documents

## Slide 22: Remaining work LLRF
  There are 4 LLRF racks, one for each 4 SSAs
  Field wiring downstairs complete – need check out
  Upstairs ~ 3 Person-Days to finish
SSA CM 1; 1 – 4
Complete
SSA CM 1; 5 - 8
Complete
SSA CM 2; 1 – 4
Need RTD & Stepper
SSA CM 2; 5 - 8
Need Stepper

## Slide 23: Remaining work JLab Cryo
  Field fit joints from cryo-can to transferline in progress
  Complete leak check of system
  Diode terminations; cables are in place (5 out of 6)
    34 diodes need to be terminated – Terminal block + 6 pcs x 32 pins
    All channels are in EPICS with proper nomenclature
  There are 3 Coriolis flow meters
    3 pcs x 9 wires need terminated in cryo-can + 3 cables to ethernet controllers mounted in the tunnel
Diode
Terminations
Resistors are
to simulate T

## Slide 24: Remaining work JLab Cryo
  3 EV; JT type – reusing existing channels rerouting existing cables
  One each EV and other flowmeter need to be tied into EPICS
  Need control screens for ease of operation
  Installation of Hot Wire Anemometer
  Complete leak check of system
  Welding and leak checking – 2 person weeks
  Hardware connections + software – 4 person weeks

## Slide 25: Remaining work Mics.
  Terminate SC magnet voltage taps in rack – 1 day
  Connect and test Fluxgate Magnetometers – couple days
    Compact Rio is setup, it with controllers will be on a cart & used only for degaussing
  Check out the Deca Rad system – 2 days
  Finish up vacuum control rack and HV cables – 1 week
  Finish bridging section clean room

## Slide 26: Summary
Work control documents are same as CEBAF
All waveguide delivered, installation complete
Cryo Can delivered and being connected
  Last piece of the puzzle
File & application servers operational, build complete
All of the cable and connectors have been delivered & installed
  Cryo cables terminated in racks
  RF cables installed and calibrated
  A few cables including cryo-can instrumentation remaining
LLRF & SSA commissioning is on-going
Positioned well for cooldown in January
  Few man-weeks of work remaining for Jlab cryo-can controls
  SRF support for CM connection to cryo-can and inter-module beam tube