## Slide 1
Integration to PID
Physics
Detectors
Machine
Community
Two basic forms of “integration”:
Deliverable measurements PID  Tracking
Space and material considerations
RICH
TOF
mRICH
DIRC
PID
Tracking

## Slide 2
FRANCOIS-XAVIER GIROD
SPENCER KLEIN
XUAN LI
XIN DONG
Håkan WENNLÖF
ASTRID MORREALE
Rey Cruz-Torres
RACHID NOUICER
Many presentations.
Similar themes:
Si as a vertex
All-Silicon

## Slide 3: From the PID Viewpoint
Cylindrical defining volume
+/- ~ 1.2 meters
Outer radius:
  ~20 cm for hybrid
  ~50 cm for All-silicon
LAURA GONELLA

## Slide 4: PID ~ Velocity
p = mgb E = mg velocity(b) measurement yields mass.
Direct measurement:
  Record signal time at multiple locations, calculate v.
  “Fast” detector = low transit time spread (most easily achieved at small transit time)
Velocity-dependent interaction(s) with detector:
  Specific Ionization (aka )
  Cherenkov Radiation:
    qC measured wrt track direction.
    Thus dependent upon deliverables from tracking
  Bremsstrahlung:
  Transition Radiation:
eID mechanisms
TOF & can be deliverables of tracking system

## Slide 5
Basic Coupling of Tracking to PID
Cherenkov Angle (qC):
  Centroid from n
  dqc from:
    Nphoto-electrons
    chromaticity (aka dispersion; n(l))
    pixel size
    optical aberration (aka “emission”)
    magnetic field
    Track Pointing
  dp from tracking
Time Of Flight (TOF):
  Calculates m2 using p, b.
  dm from:
    dt
    dp
    dL
Dominant
If dt is tiny, dL is significant
NOTE: Where PID is excellent, it can contribute to momentum resolution.
Dominant Term
Strongly reliant upon tracking
Weakly reliant upon tracking…discussed in backup slides
NOTE:
N independent TOF measurements,
Challenges exist to make independent
  Common clock
  Etc…

## Slide 6: dE/dx w/ hybrid detector
Dots = STAR
Boxes = sPHENIX Test Beam
Specific ionization is a function of b.
PID bands vs momentum.
Crossing ~b = 0.7 (TOF can help)
Limitations from Landau tails:
  Primary ionization Statistics = Poisson
  Total ionization = Landau
  Traditional Approach:
    Lots of gas & lots of samples
  Forward thinking approaches:
    Gas choice to minimize Landau
    “Cluster Counting”
Goal: Achieve dE/dx performance with least gas volume

## Slide 7
Trivial (just math) Calculation:
  Assume photon measured perfectly.
  Tracker makes fixed error a.
  dqC depends on both and a!
Requirements depend upon performance
  Parameterizations:
    Wide range of assumptions possible.
  Full Monte Carlo
    Robust and Reliable results.
Ring Imaging Cherenkov (RICH) vs. Track Direction
We present both parameterizations and full Monte Carlo calculations

## Slide 8
High Momentum GEM RICH
1m of CF4 radiator at 1.003 bar (slightly overpressure)
CsI Photocathode on top GEM
Mirror in deep UV -> MgF2 coating
Single Photon Capability ->quintuple GEM stack with APV25-SRS
Particles ~perpendicularly incident on spherical mirror, focused onto a GEM stack directly
According to the parameterization, tracking is leading error contribution is worse than ~7 mrad, becomes negligible resolution factor around 2 mrad. Between 2 and 7mrad , more detailed investigation is required.
Plot shows viscerally the effect.
More detailed simulation required.
OK
hmm
Performance from Parameterized Test Beam Data
p
p
NOTE: High chromaticity = loose requirement on tracking

## Slide 9
dRICH
Cherenkov radiator
○ Refractive Index
n = 1.02 (aerogel) 1.0008 (C2F6)
○ Length of the radiator
L = 4 cm (aerogel) , 160 cm (C2F6)
Mirrors
Photon Detector
3 mm pixel size; 200-500 nm MAPMT
Particle Generation
Originate from the vertex
Exquisite detail in simulation.
AI-based optimization.
Uses constant external angular resolution assumption.
Similar external constraint assumed as deduced for separate gas & aerogel RICHes
Tracking maybe tolerates ~2X sim (0.5 – 1.0 mrad)
NOTE: PMTs to the side makes optical aberration (“Emission”) dominant

## Slide 10
mRICH
Cherenkov radiator
(Aerogel) refractive index, n = 1.03
Length of the radiator, L = 3cm
lens with focal length, 𝑓= 6”
Photon Detector
3 mm pixel size
Tracking influence minor
Tracking influence major
NOTE: Similar performance to aerogel in dRICH

## Slide 11
DIRC
11
:
Ext. Requirements
Generic reference design: 1m barrel radius, 16 sectors
176 bars: synthetic fused silica,17mm (T) ´ 32mm (W) ´ 4200mm (L)
Photo sensors: MCP-PMTs -3x3mm2 pixels
In the absence of multiple scattering scorrelated = strack.
Secondary tracker (behind DIRC) may be beneficial both at low momentum and at high momentum (possible impact not studied yet).
All non-tracking factors sum to ~0.8 mrad, tracking must not exceed this (~ 0.5 mrad)

## Slide 12: Transition Radiation Detector (TRD)
“Threshold detector” (TR X-rays produced or not)
Pmin ~ 1 GeV (rejection ~constant at higher momentum)
1: Make TR photons:
Index, n, change
X-ray regime
Film, foam, fleece
2: Layered to get more photons
3: ID TR photons by energy/pos
4: Reject Pions
Balance between X0 and rejection

## Slide 13: Additional Integration Points
Gas RICH at highest momentum (hadron arm) does not enjoy an over-abundance of Cherenkov photons (love to have lots of Zed).
Angular resolutions required for Cherenkov-based PID detectors must be evaluated WITHIN the radiator:
  Services at Zed > 1.2 may scatter particle before hadron arm Cherenkov.
  DIRC bars scatter particles on their own (17 mm quartz)
  Must evaluate needs of a tracking layer behind Cherenkov devices.
Cherenkov detectors require photo-sensors
  Many technologies are sensitive to field strength and orientation.
  1 Tesla at the sensor is fine.
  Easy to achieve with central field = 1.5 Tesla.
  Very challenging but seems doable if central field = 3 Tesla…trade off with optical aberration as you move sensors back.
  SiPM sensors operate fine in the field, but radiation damage can lead to high rate of “false photons”
Need to minimize bending of tracks within the RICH.
Zed
Is this less than 1 mrad scattering

## Slide 14
Cherenkov Pointing:
  Parameterizations of gas Cherenkov indicate pointing required at 0.5-1.0 mrad level while inside the radiator.
  Calculations of aerogel devices indicate 0.5 -1.0 mrad level.
  Calculations of DIRC indicate 0.5 mrad level or better.
TOF:
  Path length can influence PID resolution at the best dt.
  Non-trivial influence of “material budget distribution”.
  Need deeper understanding of this issue.
PID requirements toughening with time:
  SIDIS group shown on right. Jets want even more.
  PID requirements are not symmetric in eta, driven by 0.8  1.0
  Is a symmetric barrel optimal at our asymmetric collider?
Summary
Old
New
Meta-materials: Game changer,
But we can’t assume these will work:

## Slide 15
BACKUP SLIDES

## Slide 16: Superb e/p Separation
Focused Gas RICH
Low index gas pushed the e/p out to 15 GeV or beyond.
Compromise parameter is Zed length:
  1-1.5 meters in Zed
  Under study:
    Move Zvertex toward –h exacerbating the space issue.
eRD6 RICH
prototype
eRD14 “dual RICH”
simulation
Notice:
eID needs met w threshold device
Optics maybe unnecessary

## Slide 17
psTOF
17
TOF needs a measurement of p and L
As dt➝5 ps, requires precision path length L ( c*dt~1.5 mm…tracking must exceed this.)
Multiple scattering might be the biggest worry, need to ensure low material budget and/or enough tracking layers to catch scattering.
Depends partly upon the material that does not provide position measurement:
Several technologies reviewed, no one chosen yet
AC-LGADs could provide a very good tracking layer (~100 um) + timing (20 ps)
External start time provided by forward detectors could be helpful
Study of Self-timing (Internal) using tracks

## Slide 19: More Compact Cherenkov eID Detectors
Close to 4 GeV.
Effective enough alone?
Complementary detector?
  TRD?
  dE/dx?
Modular RICH:
Fresnel-focused compact RICH
eRD14 “mRICH”
eRD14 “hpDIRC”
Useful at the lowest momenta.
Performance limited by Mult. Scattering.
Detailed simulations req’d and underway
Expectation:
  3-4s @ 1 GeV