## Slide 1: Capabilities from the Physical Sciences – Radiation Detection Selected Radiation Detection and Imaging Developments for Medical Applications A Perspective from an Experimental Nuclear Physicist
Kai Vetter
Emily Frame, Justin Ellin, Josh Cates, et al.
Department of Nuclear Engineering, UC Berkeley
Applied Nuclear Physics, LBNL

## Slide 2: Evolution of Gamma-Ray Detection in Nuclear Physics
Last 120+ years have seen enormous and still ongoing advances in nuclear instrumentation, radiation detection, and medical imaging;
Advances in gamma-ray detection for nuclear physics studies
GRETA
Gamma-Ray Energy Tracking Array

## Slide 3: GRETA & Gamma-Ray Tracking
GRETA
Gamma-Ray Energy Tracking Array
Complete shell of 120 36-fold segmented High-Purity Germanium (HPGe) detectors
Combination of 2D segmentation and pulse-shape analysis enables the reconstruction of individual gamma-ray interactions and gamma-ray tracking of up to 30 gamma-rays across the array simultaneously:
  High efficiency – ~ 60% at 1.33MeV
  High energy resolution – 0.2%
  High position resolution - ~2 mm (FHMW)
Currently under construction at LBNL to become important instrument of nuclear physics studies at the Facility for Rare Ion Beams (FRIB) at MSU.

## Slide 4: GRETA & Gamma-Ray Tracking
GRETA
Gamma-Ray Energy Tracking Array
Complete shell of 120 36-fold segmented High-Purity Germanium (HPGe) detectors
Combination of 2D segmentation and pulse-shape analysis enables the reconstruction of individual gamma-ray interactions and gamma-ray tracking of up to 30 gamma-rays across the array simultaneously:
  High efficiency – ~ 60% at 1.33MeV
  High energy resolution – 0.2%
  High position resolution - ~2 mm (FHMW)
Currently under construction at LBNL to become important instrument of nuclear physics studies at the Facility for Rare Ion Beams (FRIB) at MSU.

## Slide 5: Gamma-Ray Tracking & Imaging
Utilize gamma-ray tracking to determine location of gamma-ray source;
To simplify position determination, replace coaxial with planar detectors: Double-sided strip detectors (DSSD) built of HPGe, HPSi, or Si(Li) (or CdTe or CdZnTe):
Compton imaging
  Kinetics of Compton scattering provides incident direction (in 3D in near field)
  Gamma-ray imaging for ~150keV – 8MeV
  No collimator required → >1% efficiency
  Design to optimize angular resolution vs. efficiency → 1-2 deg. angular resolution

## Slide 6: Gamma-Ray Tracking & Imaging
Utilize gamma-ray tracking to determine location of gamma-ray source;
To simplify position determination, replace coaxial with planar detectors: Double-sided strip detectors (DSSD) built of HPGe, HPSi, or Si(Li) (or CdTe or CdZnTe):
Compton imaging
  Kinetics of Compton scattering provides incident direction (in 3D in near field)
  Gamma-ray imaging for ~150keV – 8MeV
  No collimator required → >1% efficiency
  Design to optimize angular resolution vs. efficiency → 1-2 deg. angular resolution
Si
Ge
Dimensions (using 2+2 detectors):
Si(Li):80x80x20mm3
HPGe: 80x80x30mm3

## Slide 7: Gamma-Ray Tracking & Imaging
Compton imaging
  Kinetics of Compton scattering provides incident direction (in 3D in near field)
  Gamma-ray imaging for ~150keV – 8MeV
  No collimator required → >1% efficiency
  Design to optimize angular resolution vs. efficiency → 1-2 deg. angular resolution
  HPGe + Si(Li) DSSD detectors in two cryostats & digital DAQ
  Si(Li): 32+32 strips w/ 2 mm pitch size; 10 mm thickness; 1.9 keV at 60 keV
  HPGe: 37+37 strips w/ 2 mm pitch size; 15 mm thickness; 1.7 keV at 60 keV

## Slide 8: Gamma-Ray Tracking & Imaging
2 spherical 113Sn sources (391 keV):
Diameter: Each 4 mm
Distance: 10 mm
Compton imaging
  Kinetics of Compton scattering provides incident direction (in 3D in near field)
  Gamma-ray imaging for ~150keV – 8MeV
  No collimator required → >1% efficiency
  Design to optimize angular resolution vs. efficiency → 1-2 deg. angular resolution
  HPGe + Si(Li) DSSD detectors in two cryostats & digital DAQ
  Si(Li): 32+32 strips w/ 2 mm pitch size; 10 mm thickness; 1.9 keV at 60 keV
  HPGe: 37+37 strips w/ 2 mm pitch size; 15 mm thickness; 1.7 keV at 60 keV

## Slide 9: Gamma-Ray Tracking & Compton + Coded Aperture Imaging
Compton imaging
  Kinetics of Compton scattering provides incident direction (in 3D in near field)
  Gamma-ray imaging for ~150keV – 8MeV
  No collimator required → >1% efficiency
  Design to optimize angular resolution vs. efficiency → 1-2 deg. angular resolution
Coded aperture
  Extension of multi-pinhole with overlapping projections requiring image reconstruction
  Gamma-ray imaging for <300-400 keV
  50% opaque → >10% efficiency
  Angular resolution → <1 deg.

## Slide 10: Example: Verification of ion-cancer therapy
Advances in radiation cancer therapy utilizing high Linear Energy Transfer (LET) and high Relative Biological Effectiveness (RBE) promise high efficacy in the curing of cancer, i.e.
  Targeted alpha particle therapy (TAT) and external ion beam cancer therapy (IBCT)
Since they can provide a well-localized and effective treatment, one has to ensure that the treatment is well localized within the cancer sparing the healthy tissue.
Conventional gamma-ray imaging modalities are insufficient in realizing the full potential since they require gamma-ray imaging capabilities over a broad range of energies.
TAT (e.g. Ac-225): E𝛾~200-500 keV and <1𝜇Ci and <1 mm in small animals?
IBCT (e.g. protons): E𝛾 >3MeV and 1 mm range detection in-situ within one spot (e.g. 70ms)?
Notes: TAT & HIT – but challenges in localization of dose!?
Trust but Verify…
Other imaging modalities – resolution and sensitivity
TAT: ~1uCi or less with ~mm or better resolution in small animals
HIT: ~mm range detection in-situ within one spot (e.g. 70ms)
Advances in high LET radiation cancer therapy promise high efficacy in the curing of cancer;

## Slide 11: The Promise of Targeted Alpha-Particle Therapy with 225Ac
68Ga-PSMA-11 PET/CT scans of patient with metastatic prostate cancer undergoing targeted alpha-particle therapy with 225Ac-PSMA-617. (A) tumor spread before treatment, (B) two months after the third cycle of 225Ac-PSMA-617, and (C) two months after one additional cycle.
1Kratochwil, Clemens, et al. "225Ac-PSMA-617 for PSMA-targeted α-radiation therapy of metastatic castration-resistant prostate cancer." Journal of Nuclear Medicine
Notes: Prostate-specific membrane antigen (PSMA) is a transmembrane protein that is anchored in the cell membrane of prostate cancer epithelial cells.1 PSMA has been used as a phenotypic biomarker with positron emission tomography (PET) imaging approaches.2
PSMA can be overexpressed in metastatic prostate cancer relative to normal tissue and is present in over 80% of men with prostate cancer.

## Slide 12: 225Ac Decay Characteristics and Daughter Redistribution
225Ac
day
221Fr
min
217At
ms
213Bi
min
209Tl
min
209Pb
hr
213Po
s
209Bi
Stable
MeV
MeV
MeV
MeV
MeV
How can we ensure accuracy and effectiveness?
Due to a nuclear recoil effect and different chemical properties, are alpha-emitting daughters breaking free from radiopharmaceutical preparation?
221Fr
213Bi
225Ac gamma-ray energy spectrum
Use gamma ray imaging to understand kinetics and to verify treatment!

## Slide 13: Verification of Targeted Alpha Therapy
Need for high sensitivity and good spatial resolution at 100’s of keV:
  Range → 1mm spatial resolution
  High LET → ~18kBq or 0.5𝜇Ci for mice → >1% efficiency
Coded aperture
  Extension of multi-pinhole with overlapping projections requiring image reconstruction
  Gamma-ray imaging for <300-400 keV
  50% opaque → >10% efficiency
  Angular resolution → <1 deg.
  Detectors
  2 x planar HPGe DSSDs
  76 x 76 x 15-mm3 volume
  38 x 38 strips
  2-mm strip pitch
Tungsten Mask
  Pseudorandom array
  64 x 64 square elements of 2-mm size and thickness
  Converging elements to improve image contrast

## Slide 14: Coded Aperture Tomographic Image Reconstruction of an 225Ac-filled Phantom
Experimental setup of phantom measurement in the coded aperture mode.
Max. Likelihood Expectation Maximization tomographic image reconstruction at the 218-keV line of 221Fr.
3.7 in total
10-mm separation between spheres
Verification of Targeted Alpha Therapy
~1mm spatial resolution with >1% efficiency at 218 keV

## Slide 15: Coded Aperture Tomographic Image Reconstruction of an 225Ac injected mouse (euthanized 1 hour after injection, measurement 3 days after)
Experimental setup of phantom measurement in the coded aperture mode.
Max. Likelihood Expectation Maximization tomographic image reconstruction at the 218-keV line of 221Fr.
Verification of Targeted Alpha Therapy
0.8 activity during measurement
221Fr accumulation in bladder

## Slide 16: Compton Imaging Tomographic Reconstruction of an 225Ac injected mouse (euthanized 1 hour after injection, measurement 4 days after)
Experimental setup of phantom measurement in the coded aperture mode.
Max. Likelihood Expectation Maximization tomographic image reconstruction at the 440-keV line of 213Bi.
Verification of Targeted Alpha Therapy
0.8 activity during measurement
213Bi accumulation in bladder

## Slide 17: Compton Imaging Tomographic Reconstruction of an 225Ac-filled Phantom 2mm strip pitch w/ 2 HPGe DSSDs (at 440 keV)
Experimental setup of phantom measurement in the Compton mode.
3D MLEM tomographic image reconstruction at the 440-keV line of 213Bi.
1.6 in total
10-mm separation between spheres
Histogram of an angular resolution metric for the Dual-Modality Imager. This metric was determined experimentally from a Cs-137 (662 keV) point source measurement.
FWHM(x,y) 6 mm
2-mm strip pitch
Verification of Targeted Alpha Therapy
~6mm spatial resolution with >1% efficiency at 440 keV

## Slide 18
3D MLEM image reconstruction of three simulated 440-keV point sources, each with a 10-mm separation distance, at a distance of 30 mm from an imager with 0.5-mm strip pitch.
0.5-mm strip pitch
0.5-mm position resolution
30-mm depth
Promise to obtain ~mm spatial resolution with >1% sensitivity at 440keV
Verification of Targeted Alpha Therapy
Angular resolution metric for an imager with 0.5-mm strip pitch. This metric was determined from simulated 440-keV point source data.
Compton Imaging Tomographic Reconstruction of an 225Ac-filled Phantom 0.5 mm strip pitch w/ 2 HPGe DSSDs (at 440 keV)
Under development: 80x80x15mm3 HPGe DSSD with 0.5 mm pitch and <0.3% energy resolution at >2 Mcps with benefits for nuclear physics and medical imaging studies.

## Slide 19: The promise (and challenge) of ion-beam cancer therapy
High LET and high RBE right until the end of the range.
External signal? Proton is absorbed  photon-based therapy.
How accurate can we estimate the range?
LET: Linear Energy Transfer
RBE: Relative Biological Effectiveness (e.g. p/ph~1.1)
Spread-Out Brag Peak (SOBP)

## Slide 20: The promise (and challenge) of ion-beam cancer therapy
Ideal
Actual
Conformal treatment is goal – trace edges of target volume/cancer;
Range uncertainty is dominating right now
  1mm is goal to be consistent with other uncertainties such as patient and beam alignment;
Use of lateral/anterior beams with smaller uncertainties/margins as compared to longitudinal uncertainties/range uncertainties;
Effectiveness of care and therapeutic endpoints are compromised due to additional margins.
Tang S, Both S, Bentefour H, Paly JJ, Tochner Z, Efstathiou J, et al. Improvement of prostate treatment by anterior proton fields. Int J Radiat Oncol Biol Phys. 2012;83: 408–418.
Comparison of treatment planning dose distributions for proton therapy treatment of prostate cancer with (a) one single anterior field and (b) two parallel­-opposed lateral fields.
Uncertainty in range requires larger margins in the treatment planning as should be required!
Impact of current range uncertainty

## Slide 21: The promise (and challenge) of ion-beam cancer therapy
Key: Looks like Bragg Peak!
(De) Excitation functions
Prompt gamma-ray imaging
Nuclear Physics!

## Slide 22: Verification of Ion-Beam Cancer Therapy
Prompt gamma-ray imaging at > 3MeV?
  Compton imaging for in-vivo verification?
    Coincidence timing challenge!
    Count rate: ~107 cps on average is already a challenge;
    However, this is compounded by beam bunching resulting in gamma-ray bursts of >108 cps requiring time coincidence windows of ~100s ps which is not feasible with current detectors.
  Coded-aperture?
    For high energies with high spatial resolution (~1mm) and high sensitivity (>1%) to enable the range determination with 108 protons in 70 ms (i.e. one spot of scan)?
2D knife-edge coded aperture system!
Notes: PET imaging:
  In principle, yes, since C-11, O-15, N-15, etc. are being produced;
  However, biological wash-out uncertainties due to long half-lives, creation of + emitters not at the end of the range, and/or associated with high Q-value;
PET imaging:
  In principle, yes, since C-11, O-15, N-15, etc. are being produced;
  However, biological wash-out uncertainties due to long half-lives, creation of 𝛽+ emitters not at the end of the range, and/or associated with high Q-value;

## Slide 23: Verification of Ion-Beam Cancer Therapy
New 2D knife-edge coded aperture system

## Slide 24: Verification of Ion-Beam Cancer Therapy
Line response – 2.614MeV in both dimensions
New 2D knife-edge coded aperture system

## Slide 25: Verification of Ion-Beam Cancer Therapy
Imaging results w/ 50 MeV protons from LBNL 88” cyclotron onto PMMA target (~20mm range)
New 2D knife-edge coded aperture system

## Slide 26: Verification of Ion-Beam Cancer Therapy
New 2D knife-edge coded aperture system
Estimated Bragg peak location at 58% falloff (5.6x1010 protons)
Per spot (~70ms)
Range uncertainty vs. # of protons
Timing reference?
  Can we reduce field-of-view with time-of-flight (TOF)?
    40mm w/ ~100 ps? Feasible!
  Can we determine range with TOF ?
    1mm w/ <5ps? Not feasible, yet.

## Slide 27: Conclusions
Nuclear physics and associated instrumentation can complement more conventional gamma-ray imaging approaches;
As example, current limits in high LET radiation cancer therapy can potentially be overcome with advanced concepts in gamma-ray imaging, specifically in the verification of high LET cancer therapy
  Hybrid Compton imagers or dual modality Compton/ coded aperture imager have the potential to provide ~mm resolution for ~100’s nCi activities in small animals;
  2D knife-edge/ coded aperture imager has the potential to provide ~mm range accuracy required for range verification in proton cancer therapy (potentially within 70ms).
Many other developments in semiconductor and scintillation detection systems (detectors and integrated readouts) will provide enhanced 3D position resolution and count rate capabilities.

## Slide 28
Thank you for your attention!

## Slide 29: Verification of Ion-Beam Cancer Therapy
New 2D knife-edge coded aperture system
Very preliminary imaging results w/ 67 MeV protons from UC Davis Crocker Lab onto PMMA target (~35mm range)
Timing reference?
  Can we reduce field-of-view with time-of-flight (TOF)?
    40mm w/ ~100 ps? Feasible!
  Can we determine range with TOF ?
    1mm w/ <5ps? Not feasible, yet.
Plastic scintillator

## Slide 30: Gamma-Ray Imaging – How ?
Spatial modulation-coded aperture
Gamma-ray optics?
  Difficult since gamma-rays don’t refract (much) but interact;
  Grazing incidence mirrors w/ depth graded multilayers or crystal diffraction to focus with 2-D focal plan detector
Gamma-ray interaction-based and energy dependent:
  Photo-electric absorption domain (>~300 keV)
    Collimators (2-D detectors)
    Spatial (2-D) or temporal (1-D) modulation
  Compton scattering domain (~300keV-8MeV)
    Compton imagers (3-D detectors)
  Pair production domain (>8 MeV)
    Pair “tracker” (3-D detectors)
  Positron-emission based (511 keV photons)
Compton imaging
r1
r2
r3
r4
source