## Slide 1
Superconducting Tunnel Junction (STJ)
Radiation Detectors
Stephan Friedrich, LLNL
Part 1: What are STJs?
Part 2: The BeEST Sterile Neutrino Search
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under Contract DE-AC52-07NA27344.
LLNL-PRES-811064

## Slide 2
Refresher: Si and Ge Semiconductor Detectors
Signal: QS = eN with N = E/ and
Noise: QN = e
 S/N:
Photon
(Energy E)
Conduction band (empty)
Valence band (full)
Energy resolution is proportional to
Cooling is required to reduce thermal excitations Nthermal.

## Slide 3
Semiconducting vs Superconducting Detectors
Signal: QS = eN with N = E/ and
Noise: QN = e
 S/N:
Photon
(Energy E)
Conduction band (empty)
Valence band (full)
Photon
(Energy E)
D
Valence band (full)
Signal: QS = eN with N = E/ and .7D
Noise: QN = e
1000x smaller gap  ~30x higher resolution
(and 1000x lower T needed)

## Slide 4
Superconducting Tunnel Junctions (STJs)
STJ Cross Section
STJ Band Diagram
Al
2
O
3
100×100 μm2
Nb
Ta Absorber
Al
Al
Ta
X-ray Photon
Si Substrate
SiO
2
Small ~1 meV energy gap  High resolution: ~few eV FWHM
Short ~µs charge lifetime  High speed: >1000 counts/s

## Slide 5
Statistical Noise in STJ Detectors
Charge Generation
Due to Reverse Tunneling
Due to Backtunneling

## Slide 6
Refresher: Electronic Noise in Si and Ge Detectors
Leakage Current (in Detector and FET)
Current (shot) noise
Thermal Noise (in Resistors and FET)
 Voltage (Johnson) noise of FET
Equivalent current noise
Current and Voltage Noise
Noise vs Frequency
FET noise and Si / Ge detector capacitance set electronic noise.
pA/√Hz
pA/Hz
[Goulding 1982]
in
en
Shaping time [µs]

## Slide 7
Electronic Noise in STJ Detectors
Leakage Current (in Detector and FET)
Current (shot) noise
Thermal Noise (in Resistors and FET)
 Voltage (Johnson) noise of FET
Equivalent current noise
Current and Voltage Noise
Noise vs Frequency
FET noise and STJ detector resistance set electronic noise.
pA/√Hz
pA/Hz
[Goulding 1982]
in
en
Shaping time [µs]
STJs vs Ge:
• Slower signal • Resistive detector

## Slide 8
Measuring Statistical and Electronic Noise
Ge Detectors
STJ Detectors
• Statistical noise  √E • Electronic noise is constant with E • Inhomogeneity causes broadening  E
Measured resolution
El. Noise
Statistical noise
le Grand et al., Appl. Phys. Lett. 73, 1295 (1998)
Bilger, Phys. Rev. 163, 238 (1967)
 = 2.98eV
F = 0.129

## Slide 9
STJ Fabrication
Few Pixels
36 Pixels
128 Pixels
(25µm)2 to (284µm)2
(208µm)2 pixels
(208µm)2 pixels
Our STJs are fabricated by photolithography at STAR Cryo in Santa Fe.

## Slide 10
Automated STJ Cooling to 0.1K
Adiabatic Demagnetization Refrigerators (ADRs) are compact, reliable, automated and commercially available.
1990s: 2010s:
Liquid N2 and He pre-cooling Pulse-tube pre-cooling (“dry”)
Single-stage ADR Two-stage ADR
1) Close heat switch
2) Apply B (lower entropy S)
3) Open heat switch (decouple T)
4) Reduce B slowly (keeping
entropy constant  reduce T)

## Slide 11
STJ Operation
Current-Voltage I(V) Curve (Textbook Version)
I(V) Curve (Real Life)
I. Giaever (1960)
dc Josephson
current (1962)
Cool to ≤ 0.1TC to suppress thermal current.
Apply B to suppress dc Josephson current.
Suppressed
Josephson
current
Thermal
leakage
current
Fiske modes

## Slide 12
STJ Electronics
• Low noise, stable biasing, dc voltage bias
• Computer-controlled
• 32 preamplifiers per board
Computer-Controlled Preamplifier (XIA)
Readout Requirements
• Low noise (en ≤ 1 nV/√Hz)
• Stable biasing (to ±few µV)
• Voltage bias (dc load line <100 Ω)
Specialized STJ preamplifiers are commercially available, too.

## Slide 13
STJ vs. Si Detectors
High resolution is important at low energies where lines are closely spaced.

## Slide 14
STJ Performance: Resolution and Linearity
 Energy resolution between ~1.5 and ~2.5 eV FWHM
 Only quadratic non-linearity
 Calibration accuracy of order ±1 meV in 1 hour
 Comb of peaks at integer multiples of 3.5 eV
Pulsed 355 nm (3.5eV) laser at 5,000 Hz
S. Friedrich et al., JLTP (2020)

## Slide 15
STJ Performance: Speed
M. Frank et al, RSI (1998)
M. Carpenter et al., JLPT (2017)
Nb-STJs (decay ≈ few µs)
Ta-STJs (decay ≈ few 10 µs)
STJs detectors can be operated at rates well above 1000 counts/s per pixel.
STJ Signal:
• Single exponential
decay constant
DSP pulse processing:
• Trapezoidal filter
• Pile-up rejection
~20,000 counts/s

## Slide 16
Part II: The BeEST Sterile Neutrino Search
Fermions Bosons
Leptons Quarks
What is Dark Matter?
Do Right-Handed (Sterile) Neutrinos Exist?
• Why are all neutrinos left-handed?
• Why do neutrinos have mass?
• What is 85% of the mass in the Universe?
• Why the matter-antimatter asymmetry?

## Slide 17
How to Find Something that
Doesn’t emit light
Doesn’t absorb light
Doesn’t interact (except through its mass)?

## Slide 18
How to Find Something that
Doesn’t emit light
Doesn’t absorb light
Doesn’t interact (except through its mass)?
From the recoil it causes!
|pLi-7| = |pn|
(Missing momentum experiment.)

## Slide 19
A Sterile Neutrino Search with STJ Detectors
Implant 7Be
at TRIUMF
Superconducting Detector
from LLNL
7Be
Si
STJ
Implant 7Be into STJ detectors.
1/2 = 53 days, Q = 861 keV

## Slide 20
Beryllium-7 Electron Capture in STJ Detectors
7Be + e- ⟶ 7Li + νe-
with νe-= 𝝨 |Uea|2νactive
Implant 7Be
at TRIUMF
Detect 7Be Decay
at LLNL
7Be
Si
7 Li
STJ
νe-
Measure electron capture decay of 7Be to 7Li.
2-body decay  Monochromatic recoil (in principle)

## Slide 21
Be-7 Electron Capture in STJs: The BeEST Experiment
7Be + e- ⟶ 7Li + νe-
with νe-= 𝝨 |Uea|2νactive + |Ues|2νsterile
Implant 7Be
at TRIUMF
Detect 7Be Decay
at LLNL
7Be
Si
7 Li
STJ
νe-
Heavy sterile neutrinos would change 7Li recoil energy.
Look for new peaks in recoil spectrum.

## Slide 22
The BeEST Sterile Neutrino Search
Implant 7Be
at TRIUMF
Detect 7Be Decay
at LLNL
High-Resolution Recoil Simulation
Pulsed Laser
(Calibration)
STJ
7Be
Si
7 Li
νe-
7Be + e- ⟶ 7Li + νe-
with νe-= 𝝨 |Uea|2νactive + |Ues|2νsterile
Calibrate STJ with pulsed laser.
Four peaks due to K- and L- capture into 7Li ground and excited state

## Slide 23
The BeEST Sterile Neutrino Search
Implant 7Be
at TRIUMF
Detect 7Be Decay
at LLNL
Pulsed Laser
(Calibration)
STJ
7Be
Si
7 Li
νe-
7Be + e- ⟶ 7Li + νe-
with νe-= 𝝨 |Uea|2νactive + |Ues|2νsterile
Data from 1 Ta-STJ Detector for 28 Days
• 4 primary peaks – 2 x K-capture, 2x L-capture
– to 7Li ground state and to 7Li*
• 4 high-energy tails
– Shake-off effects
• 2 low-energy tails
– (Partial) Auger e- energy loss
• 1 broad background
– 478 keV ‘s in substate
L/K Ratio = 0.070(7)
PRL 125, 032701 (2020)

## Slide 24
Data with Hypothetical Sterile v Signal
Exclusion Plot
10x improvement with 1 pixel.
PRL 126, 021803 (2021)
The BeEST Sterile Neutrino Search
Hypothetical
sterile neutrino signal

## Slide 25
Summary of Sterile Neutrino Searches (2021)
Tabletop BeEST Experiment
The BeEST in Context
G$ experiments
m4 [GeV]
BeEST
2021
From: Physics Reports 928, 1 (2021)

## Slide 26
Current Work: New Materials and More Pixels
STJs from Different Materials
Currently Taking Data with Ta-STJ Array
Stay tuned!
Ta-STJ Detector
Nb-STJ Detector
Al-STJ Detector
Separate BSM physics from material effects.

## Slide 27
Summary: STJs and the BeEST
Superconducting Tunnel Junctions
Beryllium-7 Electron Capture in STJs
• High energy resolution: ~1 – 5 eV
• High speed: >1000 counts/s per pixel
• No quenching
• High-sensitivity recoil measurement
• Exclusion to |Ue4|2 to ~10-4 with 1 pixel
• Currently scaling to arrays

## Slide 28: Thank You!
Robin Cantor
Ad Hall
 STJ Fab
Bill Warburton
J. Harris
 STJ Electronics
Stephan Friedrich
G.-B. Kim, I. Kim, A. Samanta, V. Lordi
BeEST Experiment
 DFT Simulations
Chris Ruiz
A. Lennarz, P. Machule, D. McKeen
7Be Implantation
Cosmology
Xavier Mougeot
EC Decay Simulations
Jose-Paolo Santos
J. Machado, M. Guerra
Shake-up/off Simulations
Cynthia Volkert
H. Hadenfeldt, J. Arlt
 TEM Imaging
Kyle Leach
S. Fretwell, C. Bray, A. Marino, C. Harris
Weak Interaction Physics
Monte-Carlo Simulations