## Slide 1: Status of HRS Detectors
Barak Schmookler

## Slide 2: Overview
Standard detector stack
Upgrades*:
  VDC: Replaced the old discriminator cards
  One straw chamber from FPP is installed in each arm
  Gas Cherenkov: New windows, new PMTs, S0 on front
  *Partially Completed

## Slide 4: VDC
Replaced discriminator cards with new MAD cards
HRS-L: gas, DAQ restored & tested
HRS-R: in progress
Install new threshold PS module for cards (Dec 2013): currently set to 3V for cosmic data runs

## Slide 5: FPP
Straw chambers installed in each arm
Gas operational in L-HRS; need clean gas line in R-HRS
All cables installed; HV distribution active
Low voltage supply for read-out modules operational (Nov 2013)
Signals checkout: in progress

## Slide 6: Gas Cherenkov/S0
Gas Cherenkov /S0 installed in both arms
Added extension to L-HRS Gas Cherenkov
New windows added; mirror alignment completed
PMTs should be installed in Jan-Feb 2014

## Slide 7: Gas Cherenkov/S0

## Slide 8: Gas Cherenkov: New Windows & Mirror Alignment
Replaced 3 out of 4 windows
New frame is half thickness of old frame; new aluminum windows are thinner
We could not install new window in back of L-HRS GC, because of fixtures attached to two mirrors
Used old window with aluminum plate instead
Placed light source a certain distance from mirrors and adjusted image located at PMT position

## Slide 9: Gas Cherenkov: New PMTs
Mirror Reflectivity tests done
Deterioration in reflectivity was observed for all mirrors tested
Reduction in relative efficiency for BURLE PMTs ~7%

## Slide 10: Gas Cherenkov: New PMTs
Will use new PMTs (ET Enterprises, model: 9390KB)
Tested 4 different types of wavelength –shifting (WLS) paint to increase efficiency
  Goal: At least 10 PE in the experiment

## Slide 11: Gas Cherenkov: New PMTs

## Slide 12: Gas Cherenkov: New PMTs
Tested four different WLS paint samples:
– EJ-289 #2 (commercially available)
– EJ-299-31A
– EJ-299-31C Evaluation samples specially developed by Eljen Technologies
– EJ-299-31E
Sample EJ-299-31E gave best improvement in number of photo-electrons
-------------------------------------------------------------------------------------------------
Setup # painted(y/n) # of PE % increase in gain
-------------------------------------------------------------------------------------------------
A no 10.8 -
B no 10.1 -
C no 6.13 -
-------------------------------------------------------------------------------------------------
A yes 14.3 32
B yes 13.8 37
C yes 8.63 41
-------------------------------------------------------------------------------------------------
Different Setups correspond to different locations of the silica disks
Results are reproducible with different PMTs and over time (~2 weeks)

## Slide 13: Gas Cherenkov: New PMTs
Tested four different WLS paint samples:
– EJ-289 #2 (commercially available)
– EJ-299-31A
– EJ-299-31C Evaluation samples specially developed by Eljen Technologies
– EJ-299-31E
Sample EJ-299-31E gave best improvement in number of photo-electrons
-------------------------------------------------------------------------------------------------
Setup # painted(y/n) # of PE % increase in gain
-------------------------------------------------------------------------------------------------
A no 10.8 -
B no 10.1 -
C no 6.13 -
-------------------------------------------------------------------------------------------------
A yes 14.3 32
B yes 13.8 37
C yes 8.63 41
-------------------------------------------------------------------------------------------------
Different Setups correspond to different locations of the silica disks
Results are reproducible with different PMTs and over time (~2 weeks)
Paint turns cloudy after 2-3 days but this does not affect QE

## Slide 14: Data Acquisition
Took data in L-HRSwith cosmic particles over last week
Triggered on coincidence between S2M and old S0
After some work: VDC, S2M, and Calorimeter are functioning well. (Need to perform gain matching)
FPP map/database is complete; but need to determine which wire in bunch was hit

## Slide 15: Summary
VDC upgrades are completed
Gas Cherenkov/S0 should be operational in Jan-Feb 2014
Work continues on FPP chambers