## Slide 1: DOE/NIH workshop: Advancing Medical Care through Discovery in the Physical Sciences 2023
Emilie Roncali1,2
Cameras and detectors for radiopharmaceutical therapy:
State-of-the-art, Challenges and Emerging Technologies
UCD Health Medical Center, Sacramento
1Radiology
UC Davis
2Biomedical Engineering
Notes: Focus on radio-theranostics and nuclear imaging

## Slide 2: Outline
Radiopharmaceutical therapy and dosimetry
The need for quantitative imaging for personalized treatment
Challenges and solutions for image-based dosimetry
Notes: My talk focuses on radio theranostics applications and not so much on EBRT but this will be discussed in the session

## Slide 3: Radiopharmaceutical Therapy: Definitions
At the interface of chemistry, radiobiology, immunology, physics, diagnostics imaging, pharmacology
Not new!
Also called…
Targeted Radionuclide Therapy (TRT)
Molecular Radiotherapy (MRT)
Locally delivered radiation therapy through biological targeting
I.M. Ariel and G.T. Pack, Cancer, 1967

## Slide 4: RPT and Targeting
Peptide Receptor Targeting
177Lu-Dotatate
177Lu
FDA-approved 2018
gNETs
Antibody targeting
90Y Ibritumomab tiuxetan
FDA-approved 2002
B-cell non-Hodgkin lymphoma
Prostate-specific Membrane Antigen
FDA-approved 2022
mCRPC
177Lu-PSMA-617
RPT now indicated in 35% new cancers (33% deaths)
β emitter, with small positron emission
Image with bremsstrahlung SPECT or PET
β emitter with 2 γ emissions
Image with SPECT
β emitter with 2 γ emissions
Image with SPECT
Notes: Radionuclide therapy can be targeted in the same manner as other pharmaceuticals
131I- Sodium Iodide 1971 with Iodine 131
- 131I Tositumomab Bexxar 2003-2014 or Y90 ibritumomab are antibody targeting. Bexxar used for B-cell lymphoma and tiuexetan (Y90) used for B non-Hodgkin lymphoma. Murine monoclonal antibody selectively binding to B-lymphocyte antigen CD20 on surface of B lymphocytes
- Ra-223 Xofigo 2013- is radium dichloride, bone-seeking Ca mimetic that targets areas with bone turnovers, such as mets. Alpha emitter 82 keV
Y-90 spheres
PSMA is an enzyme overexpressed in prostate cancer and mets castrate-resistant cancer
Lu-DOTATATE gastro entero pancreatic neuroendocrine tumors (GEP-NETs) that are positive for the hormone receptor somatostatin, including GEP-NETs in the foregut, midgut, and hindgut.
Reminder: Antigens can be proteins, peptides (amino acid chains), polysaccharides (chains of monosaccharides/simple sugars), lipids, or nucleic acids

## Slide 5: Radionuclide Therapy emitters
C. Miller, J. Rousseau, C. F. Ramogida, A. Celler, A. Rahmim and C. F. Uribe. “Implications of physics, chemistry and biology for dosimetry calculations using theranostic pairs”, Theranostics 2022
 | 153Sm | 1.95 d | β- | SPECT | 103 | 30 | 810 | 
 | 166Ho | 26.8 h | β- | MRI, Brems. | 81 | 7 | 1850 | 
 | 213Bi | 45.6 min | Β-, α | SPECT | 440 | 16.4 |  | 8.4
 | 223Ra | 11.4 d | β-,α | SPECT | 82 | 20 |  | 5.97
emitters
PET, SPECT, Bremsstrahlung
Role of theranostic imaging and dosimetry?
Notes: α decay:
Short range attractive for targeted therapy (tens of microns)
Radium-223 dichloride FDA approved for metastatic prostate cancer (bone)
Ra-223 (Xofigo)
Alpha-particle emitter (~80 µm range in tissue; 11.8 day half-life)
223RaCl – calcium analog; concentrates in rapidly proliferating bone
55kBq (1.49 µCi)/kg every 4 weeks for 6 cycles → 2 month survival improvement
Protocol to image patients by quantitative SPECT/CT after 1st and last cycle
Ac-225, At-211, Bi-213 in development
- decay
Lu-177
  Lu-177 dotatate for gastro entero pancreatic neuroendocrine tumors
  Lu-177 PSMA for metastatic prostate cancer
  Energy max = 479 keV, max range 2.7 mm, 2 useful photons for SPECT (112 keV and 208 keV)
Y-90 DOTATE or Y-90 microspheres for liver radioembolization
  Energy max =2.28 MeV, max range 11mm
  Can be imaged through Bremsstrahlung emission (SPECT) or low positron emission (PET)

## Slide 6: Quantitative nuclear imaging and dosimetry for radiotheranostics
Radionuclide therapy
177Lu- PRRT
177Lu
FDA-approved 2018
Theranostics is redesigning the nuclear medicine landscape
Uribe et al JNM 2021
Dosimetry
Ideal plan for each patient:
We see  we plan  we treat  we monitor response  we adjust and retreat
Quantitative imaging at the core of dosimetry
Technological gaps to support this paradigm create opportunities for new developments
Notes: We see= diagnostic imaging, typically SPECT or PET/CT
We plan – predict the dose in OAR and tumors if possible
We treat = we deliver the RPT in 1 cycle, typically several cycles are planned based on FDA approved molecules
We monitor response = we assess the dose in OAR and tumors. This is where imaging plays a big role, we need PET/CT or SPECT/CT again
We re-adjust the dosimetry planning to respect toxicity and tumor dose (do we know the limits?)
Real integration of imaging and therapy for cancer!

## Slide 7: Dosimetry improves patient outcome: Y-90 DOSISPHERE-01 trial and TARGET Study
Personalized dosimetry improves patient outcome
Tumor dose correlates with response
Notes: A randomized, multicenter, investigator sponsored phase II trial comparing the clinical outcomes of SIRT with TheraSphere in patients with intermediate/advanced HCC using two pre-treatment dosimetry determination methods: (1) Standard, single-compartment dosimetry (SDA); defined as a uniform distribution of absorbed dose within the perfused volume – both tumor and normal liver or (2) Personalized dosimetry (PDA); defined as multi-compartment Y-90 distribution of absorbed dose within the perfused volume that accounts for preferential blood flow into the tumor compared with normal parenchyma.
Median OS 26.6 months with personalized instead of 10.7 months standard multicompartment dosimetry
Personalized means 205 Gy to tumor and 120 Gy rest liver instead of 120 Gy global

## Slide 8: Image-based dosimetry
-50
50
-50
50
0
0
mm
Dose Point Kernel
Absorbed dose
Total activity
Absorbed dose
Radiation transport model
Total activity + anatomy (CT, MRI)
How do we quantify the total activity?
Notes: Typically calculated in average dose in shells
Dose point kernel, or specific absorbed dose point kernel, relates to the energy deposited at any distance from the source
DPK depend on electron range
CSDA= continuously slowing down approximation:
= path length that an electron would traverse when slowing down from its original energy E. to a stop, if its rate of energy loss along the track were equal to the mean rate of energy loss.

## Slide 9: Post –treatment 90Y Imaging with Total-body PET
Liver segmental dosimetry
uEXPLORER (UIH), 194 cm axial FOV
Costa et al. Journal of Nuclear Medicine 2021
275.3 Gy
275.0 Gy
gNET patient: (3.36 GBq 90Y)
30 min scan 5 hours after administration
Notes: 3 injections to right lobe
uEXPLORER 4 mm isotropic voxels, PSF on, no smoothing
Y-90 PET shows microspheres in liver segments 5,6,7, & 8 where lesions are
Segment tumor volume to build tumor and nontumoral liver dose volume histograms
Introduce Yuni’s talk on dosimetry

## Slide 10: Challenge: dose heterogeneity in liver 90Y radioembolization
Y-90 PET/CT and arterial tree
Y-90 PET/CT + CBCT + CECT
Simulation of microsphere distribution
CBCT
Heterogeneity comes from microsphere distribution in hepatic artery
Notes: Add that with Y-90 DOTATATE we have low activities which are also challenging

## Slide 11: Looking forward: How can we improve image-based dosimetry?
Activity

Dose kernel
Current clinics future technology
Improve image quantification & dosimetry
(computational)
Improve imaging systems
(hardware)
Work with Radiation Oncology
Improve technology
Design new instruments
Work with Engineering (BME, ECE…)
High resolution
High sensitivity
Fast imaging / long axial field of view
Portability
Notes:   A multidisciplinary challenge
  team effort
  TB PET, app specific PET and better collim, energy res to reject scatter in SPECT
  I frame this in the context of personalized q Dosimetry but there are other important aspects (AI , kinetic modelling, image reconstruction)
  A lot of the tech I am discussing next is for PET but some applies to SPECT

## Slide 12: High resolution imaging for multiscale dosimetry
Elgqvist J, et al.. The potential and hurdles of targeted alpha therapy - clinical trials and beyond. Front Oncol. 2014
α range ~50–70 μm
β range ~700 μm
100 μm
Ra-223 SPECT/CT Reconstructed Images
Micrometastases from ovarian cancer in peritoneal lining
Courtesy George Sgouros, Eric Frey
Several orders of magnitude differences… How can clinical scans inform on α dosimetry?
Notes: Alpha-emitter imaging is challenging. Some have deemed it impossible
Very high potency - > low administered activity ->very low count rate/noisy projections
Eric Frey quantitative recon method: models and compensates for image degrading effects
– no photon left behind– implemented in Rapid recon algorithm
Alpha SPECT challenges
Complicated decay schemes with multiple daughters
Multiple gamma and x-ray emissions
Presence of high energy photons
Small administered activities (Ra-233 =55kBq/kg x 70 kg 1 mCi vs 200 mCi for beta-emiters e.g. PSMA-617)
Need spectrometry
Micromets figure
Left: range of the α-particles can hardly reach the surrounding normal healthy cells other than possibly the mesothelium and its sub-layer. cannot reach the epithelial cells of the intestinal lining

## Slide 13: High-sensitivity and fast imaging for dosimetry
Field of view for Lu-177 PRRT SPECT/CT is neck to pelvis, or head to pelvis  2-3 bed positions
Depending on uncertainty and time point post injection (7 days) scan time can be >1h, repeated 5 times per cycle
High sensitivity or fast imaging needed
Solutions for high-sensitivity:
Collimator
Crystal NaI(Tl)
Collimator-less imaging
Direct conversion with semiconductors CdZnTe detectors
Notes: 2-3 bed positions takes time to acquire enough counts in each projection, typically at least 60 proj and 5s per project
Mention stationary scanners based on CZT

## Slide 14: High-sensitivity SPECT
Caravaca J, et al. Compton and proximity imaging of 225Ac in vivo with a CZT gamma camera: a proof of principle with simulations. IEEE TRPMS. 2022 R01 EB026331
Compton imaging
Alpha SPECT mini
L. Cai et al. An Ultrahigh Energy Resolution SPECT System for Quantitative Hyperspectral Imaging of Targeted Alpha Therapy.
JNM June 2022, 63 (supplement 2) 3305;
Notes: Proximity imaging higher sensitivity Compton can perform better than conventional SPECT
Can be used for alpha
Alpha SPECT has 24 modules spectrometers with multi readout circuitry and hyperspectral. Collimator has 96 knife-edge pinholes of 1mm diameter
Will support research as well.

## Slide 15: Targeted Alpha therapy imaging (Ac-225)
Dual modality imager Coded aperture + Compton
E. A. Frame et al.. Coded Aperture and Compton Imaging for the Development of Targeted Alpha-Particle Therapy. arXiv preprint arXiv:2212.07631 2022 DE-NA0003180
γ 218 keV
 coded aperture
γ 440 keV
 Compton imaging
Notes: 3D position-sensitive high-purity germanium (HPGe) double-
sided strip detectors (DSSDs)
planar
- HPGe crystal with an active volume of 74 74 15 mm3
- coded aperture is fabricated from 2:4-mm-thick tungsten, which ensures
roughly 90% attenuation at 250 keV. The mask consists of 64 64 square elements
with each individual element having a 2 2 mm2 face to match the detector pixel size
and limit collimation eects
- coded aperture combined with the front detector constitutes the coded aperture
imager, while the front and back detectors collectively form the Compton camera

## Slide 16: Imaging for dosimetry: how many time points are needed?
Single time point imaging using population data to estimate the organ clearance
2021
Kidney dosimetry in gNETs treated with 177Lu-DOTATATE
Mixed models demonstrate potential with 1-2 scans
2021
Not there yet for other organs…
Poor accuracy in bone marrow and tumors for 177Lu PSMA
How to acquire multiple scans in a simpler, more cost-effective way?
Notes: Madsen:
Simulations were performed using both exponential models where the rate constants and associated parameters were randomly sampled from distributions with a known mean. The actual total integrated activity for each random sample was compared with the estimated total integrated activity using the mean value of the parameters. Retrospective analysis of 90Y DOTATOC data from a clinical trial provided a comparison of actual kidney dose with the estimated kidney dose using the single time point approach.
Conclusion: When there is prior knowledge about the population averages of kinetic parameters, these results suggest that the single time point approach can be used to estimate the total integrated activity and dose with sufficient accuracy to manage radionuclide therapy. This will make personalized dosimetry much easier to perform and more available to the community
Devasia: to evaluate a novel method for joint kidney time–activity estimation based on a statistical mixed model,
A prior 10 gNETs patients with complete time–activity data,
Only 1or 2 imaging points for new patients was compared with previously proposed single-time-point methods in virtual and clinical patient data.
Conclusion: Mixed models based on a historical cohort of patients with complete time–activity data and new patients with only 1 or 2 SPECT/CT scans demonstrate less bias on average and significantly fewer outliers when estimating kidney TIA, compared with popular reduced-time point methods.
Use of mixed models allows for reduction of the imaging burden while maintaining accuracy, which is crucial for clinical implementation of dosimetry-based treatment.
Hou 2021: large error several folds in bone marrow
Although there’s a lot of work done in improving STP dosimetry what alternative do we have? SEND the patient home!

## Slide 17: At-home dosimetry: Remote radiation tracking
Wearable dosimetry
Collect real-time dosimetry for days without serial hospital visits
Requirements
Design and Initial Performance Characterization of a Personalized Remote Radiation Tracking Portable Organ Dosimetry Device (PRRT PODD) IEEE MIC 2022, R. S. Miyaoka, L. A. Pierce, R. L. Harrison, D. Q. DeWitt, H. Vesselle
Notes: Cost effective, user friendly personalized dosimetry for Lu-177 DOTATE therapy
Miniaturized system 10-15 small modules
Rugged registration
Body positioning
Portable electronics

## Slide 18: Novel detectors
Cerenkov-based detectors
Kwon et al BPEX (2019)
Cerenkov BGO detector 280 ps
Cerenkov MC-PMT detector 32 ps
Kwon et al Nature Photonics (2021)
Metascintillators
Turtos et al PMB (2019)
LYSO/BC422 76 ps
DOI and timing
Cerenkov Charge Induced TlBr detector 330 ps
Ariño-Estrada et al PMB (2019)
Applications:
PET imaging
proton gamma imaging
new SPECT detectors
Notes: Mention material design

## Slide 19: Conclusions
Quantitative nuclear medicine has made tremendous progress!
Integration of imaging and therapy is a strong driver of technological and computational innovation
Many exciting opportunities ahead of us to deliver robust quantification accuracy!
  Technology
  Standardization
  Image processing

## Slide 20: Acknowledgments
Biomedical Engineering
Gustavo Costa, Ph.D.
Carlos Ruvalcaba, Ph.D
Amirtaha Taebi, Ph.D.
Carlotta Trigila , Ph.D.
Gerard Ariño-Estrada, Ph.D.
Sun-Il Kwon, Ph.D.
Simon Cherry, Ph.D.
MIPET group
Julie Sutcliffe, Ph.D.
Radiation Oncology
Stanley Benedict, Ph.D.
R01 EB027130
R21 CA237686
DDCF
Radiology
Ramsey Badawi, Ph.D.
Denise Caudle, CNMT
Cameron Foster, M.D.
Heather Hunt, CNMT
Negar Omidvari, Ph.D.
Rex Pillai, M.D.
Benjamin Spencer, Ph.D.
Catherine Vu, M.D.