## Slide 1: Image Reconstruction and Processing for Alpha Emitters Yong Du. Ph.D. Department of Radiology, Johns Hopkins University

## Slide 2: Conflict of Interest Disclosure
Under a licensing agreement between the Radiopharmaceutical Imaging and Dosimetry, LLC and the Johns Hopkins University, Yong Du is entitled to a share of royalty received by the University on sales of iterative reconstruction software used to obtain some results in this presentation.
This interest has been disclosed and is being managed by the Johns Hopkins University in accordance with its conflict-of-interest policies.

## Slide 3: Radiopharmaceutical Therapy (RPT)
Radiopharmaceutical therapy (RPT)
  Beta: 131I, 177Lu-DOTATATE, 177Lu-PSMA, 90Y …
    103 to 104 tracks to kill cell
    DNA damage is repaired
  Alpha: 225Ac-PSMA, 223Ra-Xafigo, …
    2 to 3 tracks kill cell
    Irreparable DNA damage
Optimal use of RPT requires patient-specific dosimetry
Sgouros, G., et al. Nat Rev Drug Discov 19, 589–608 (2020).

## Slide 4: Dosimetry for RPT
Activity A(t) (MBq)
Time t (hours to days)
=
Time integrated activity (TIA)

## Slide 5: Quantitative SPECT Activity Estimation
CT
0 hr
4 hr
24 hr
72 hr
144 hr
SPECT
SPECT Reconstruction & Convert to Activity
Sum
Activity in
VOI
SPECT
Proj.
CT
Total
Liver
Activity
Measurement
Estimate

## Slide 6: SPECT TIA Estimation
Time-integrated Activity
CT
SPECT
Proj.
Curve Fitting
SPECT
Activity
Estimation
0 hr
4 hr
24 hr
72 hr
144 hr

## Slide 7: Challenges with SPECT Imaging of Alpha-Emitter
Complicated decay schemes with multiple daughters
Daughter with long half live could have different distribution
Multiple gamma and x-ray emissions with low abundances
Presence of high energy photons
Small administered activities

## Slide 8: Alpha Emitter Example: Ac-225
Measured spectrum of 225Ac point source in air (blue line) and 225Ac activity uniformly distributed inside a 6 cm diameter sphere in a 22 cm diameter water cylinder (yellow line). The important photopeaks and their abundance are marked with dotted line.

## Slide 9: Ra-223

## Slide 10: Th-227
Imageable Photons
E ≥ 50 keV
Abundance ≥ 10%
Energy (keV) | Abundance (%) | RN
81.4 | 15.7 | Ra-223
84.1 | 25.9 | Ra-223
236.0 | 12.9 | Th-227
269.5 | 13.9 | Ra-223
271.2 | 10.8 | Rn-219
351.1 | 12.9 | Bi-211

## Slide 11: Solutions for Imaging Alpha Emitters
Optimize imaging protocol
Developed advanced reconstruction algorithms
  Compensation for imaging physics
  Multiple radionuclide reconstruction
Developed next generation SPECT cameras
  Semiconduction detector  energy resolution
  New collimator  sensitivity and spatial resolution

## Slide 12: Model-Based Reconstruction with Crosstalk Compensation
Ra-223 Projection image Ra win
Ra-223 Activity Distribution
Recon I (A,D,S)*
Ra win
Th-227 Projection image Th win
Recon I (A,D,S)*
Th win
Th-227 Activity Distribution
Project (A,D,S)
Th win
Project (A,D,S)
Ra win
Ra-223 Crosstalk Estimate in Th win
Th-227 Crosstalk Estimate in Ra win
Recon II (A,D,S,Crosstalk)*
Ra win
Recon II (A,D,S,Crosstalk)*
Th win
Final Th-227 Image
Recon III (A,D,S,Crosstalk)
Th win
Final Ra-223 Image
Recon III (A,D,S,Crosstalk)
Ra win
Yes
No
Converge?
*A: Attenuation, D: Collimator-Detector response, S: Scatter

## Slide 13: Simultaneous Reconstruction
227Th images
223Ra images
Simultaneous multiple energy window multiple radionuclide OS-EM
236 keV energy window projections
154 keV energy window projections
269 keV energy window projections

## Slide 14: Th-227/Ra-223 Phantom Study
Blue spheres contain 223Ra.
Clear spheres contain 227Th.
Background filled with water
In collaboration with Bayer
Siemens Symbia SPECT/CT system, ME collimator

## Slide 15: Day 1
227Th
227Th
227Th
227Th
227Th
227Th
223Ra
223Ra
223Ra
223Ra
223Ra
227Th images
223Ra images
223Ra
reference
ME collimator, 10 seconds acquisition, 2 iterations 12 subsets, no post recon filtering

## Slide 16
227Th
227Th
227Th
227Th
227Th
227Th
223Ra
223Ra
223Ra
223Ra
223Ra
227Th images
223Ra images
223Ra
reference
Day 14
ME collimator, 10 seconds acquisition, 2 iterations 12 subsets, no post recon filtering

## Slide 17: Ra-223 Reconstructed Images
5 coronal slices, 144 h, 5 iterations, 12 subsets per iteration, logarithmic grayscale
Simulation study using XCAT phantom and SIMIND Monte Carlo simulation program

## Slide 18: Animal Studies
221Fr
24 h
72 h
144 h
24 h
72 h
144 h
Francium-221
Bismuth-213
Fused SPECT/CT images of rabbit with a VX2 hepatic tumor at 24, 72 and 144-hours post-injection of 50 uCi [225Ac]Ac-DOTA-TDA/Lipiodol emulsion. Left Panel represents the Francium-221 energy window. Right panels Represent the Bismuth-213 energy window. (Du, et al, PLoS ONE 17(1) e0261982)
Image acquired on a Siemens Symbia T16 SPECT/CT

## Slide 19: Ac-225 Phantom Experiment
22 cm
  Goal: Assess Accuracy
  6 cm diameter sphere
  Cold background
  4.76 Mbq of Ac-225 in sphere (at equilibrium)
  22 cm diameter cylindrical phantom
  60 views at 150 s per view (75 min)
  Siemens Symbia Scanner
  HEGP Collimator
  Fr-221 and Bi-213 energy windows

## Slide 20: Ac-225 Phantom Experiment: Accuracy
Reconstructed images (20 iterations)
Radionuclide | Error (%)
Fr-221 | -8%
Bi-213 | 1%
Fr-221 Bi-213

## Slide 21: Ac-225: Precision of Activity Estimates
Fr-221
Bi-213
0.60 MBq
0.24 MBq
0.12 MBq
6 iterations, 12 subsets
30s/view, 120 projections, 50 noise realizations
Activity (MBq) | Fr-221 | Bi-213
0.60 | 1.84% | 1.52%
0.24 | 2.85% | 2.39%
0.12 | 5.97% | 5.38%
0.02 | 16.24% | 12.16%
Coefficient of Variation
0.02 MBq

## Slide 22: Conclusion
SPECT imaging of alpha emitting radionuclides
  Feasible with current clinical systems
  Requires careful attention to modeling the imaging physics
  Accuracy ~10% for organ-sized object
  Can provide information about distribution of daughter radionuclides

## Slide 23: Acknowledgements
R01 EB031023
U01 EB031798
U01 CA140204
R01 CA239041
Eric Frey, Ph.D.
George Sgouros, Ph.D.
Martin Pomper, M.D. , Ph.D.
Ling-Jian Meng, Ph.D
Jessie Nedrow, Ph.D.
Rebecca Krimins, Ph.D.
Rob Hobbs, Ph.D.
Dara Kraitchman, Ph.D.
Bin He, Ph.D.
Michael Ghaly, Ph.D.
Stavroula Sofou, Ph.D.
Remco Bastiaannet, Ph.D.
Ioanna Liatsou , Ph.D.
Junyu Chen, Ph.D.
…

## Slide 24
Thank you!