## Slide 1: Characterizing integrating detector PMTs for the MOLLER Experiment
Sagar Regmi
Idaho State University
HUGS 2023

## Slide 2: Outline
1
sagarregmi@isu.edu
Parity Violating Asymmetry
MOLLER Introduction
PMT and Systematic Error
Non-Linearity measurement setup
Calculations, Observed Graphs
Summary and Future work

## Slide 3: Parity Violating Asymmetry
sagarregmi@isu.edu
2
Keeping spatial vector unchanged with spin reversed
The scattering rate under parity transformation is invariant in all interactions except the weak interaction.
The asymmetry can be calculated as,
where and are scattering cross section for right- and left-hand helicity window.
Under Parity Transformation there is inversion of spatial vector , no change in
Scattering of polarized incident electron (Spin S) and unpolarized target

## Slide 4: MOLLER INTRODUCTION
sagarregmi@isu.edu
3
Measurement Of Lepton-Lepton Electroweak Reaction.
Measures weak charge of electron using Parity Violating electron Scattering (PVeS); measurement precision allows for significant test of electroweak physics and search for physics beyond the standard model
The parity violating asymmetry in the scattering is:
The weak charge of electron is:
(electroweak theory prediction at tree level)
Where is the electroweak mixing angle..
Prediction
APV = 35.6 ppb
δ(QWe) = ± 2.1 % (stat) ± 1.0 % (syst)
δ(APV) = 0.73 parts per billion, δ( ) = ±0.00028
signal rate: 135 GHz run time: 8200 hours

## Slide 5: MOLLER Intro… and PMT
sagarregmi@isu.edu
4
The photomultiplier tube (PMT) converts a continuous stream of photons into a steady anode current during a helicity window, which is then converted to a voltage signal using a preamplifier and digitized by an 18-bit ADC for integration over the same window.
PMT
Components
CEBAF beam: ~90% longitudinally polarized, 11 GeV, Helicity reversal @ 960 Hz
Toroidal Spectrometer: Upstream and downstream, focuses the full momentum range of scattered particles ranging 2-8 GeV.
Tracking Chambers
Integrating detector: Main int det, Shower-max detector, Pion background detector
Aux detectors: SAMs, LAMs, Scanners

## Slide 6: Systematic Error
sagarregmi@isu.edu
5
Summary of projected fractional statistical and systematic errors on the parity-violating asymmetry.
PMT non-linearity is an important systematic error and needs to be understood well before experiment
There are many source of systematic error like kinematic factor, Beam(second moment, polarization, Beam intensity, radioactive corrections, background radiations and PMT non-linearity
The non-linearity of the detector is mainly caused by the PMT gain's non-linear response to its cathode current.

## Slide 7: Non-Linearity Measurement setup
sagarregmi@isu.edu
6
Fig : CAD view of bench-test setup used for PMT non-linearity calibration. This assembly is kept in a light-tight box while taking data.
Non-Flashing LED provides overall Light Level (Cathode Current) and Flashing LED operated at helicity reversal frequency 960Hz provides small LED asymmetry.
The filter transmission settings are: 1, 10, 25, 40, 50, 63, 79, and 100 %. These were pseudo randomly arranged around the wheel
The synchronized square wave from an arbitrary function generator is used to control both the intensity and frequency of the flashing LED, as well as the DAQ frequency.
The DAQ records ON and OFF state of LED.
The CODA (CEBAF Online Data Acquisition) software package is used for data collection.

## Slide 8: Calculations
sagarregmi@isu.edu
7
The first order non-linearity of the measured PMT signal asymmetry is given by
≈ Atrue(1 + βNavg), With
Fits to ALED vs. N avg or Vadc plot give Atrue and Atrueβ and hence the non-linearity.

## Slide 9: Results and Observed Graphs
sagarregmi@isu.edu
8
Plot: Raw, Pedestal, Corrected and Asymmetry Distributions(7.0 nA LL)
Plot: Filterwise Asymmetry Distributions for 1st filter cycle(7.0 nA LL)

## Slide 10: Summary and Future works
sagarregmi@isu.edu
9
Plot: ALED vs VADC plots and non-linearity calculation(7.0 nA LL)
We developed a bench-test setup to characterize the PMT non-linearity, which mimicked the same technique of asymmetry measurement as in the precision electroweak experiments like PREX-2 and CREX.
We have different pmt base for MOLLER compared to PREX/CREX, in which the preamplifier is embedded in it. We are also using different ADC.
We are also trying to use a different software package then CODA to record the events.
However, the technique of measuring non-linearity has not changed with the new setup.
Other Works: Designing of Shower Max Detectors, Radiation hardness Testing

## Slide 11
sagarregmi@isu.edu
10
Acknowledgement
Prof. Dr. Dustin E. McNulty, Chair of Department of Physics, Idaho State University
Dr. Devi L. Adhikari, Research Associate, Virginia Tech
Sudip Bhattarai, Graduate Student, Idaho State University
MOLLER collaborators
HUGS Program 2023 organizers and participants

## Slide 12: Thank you
sagarregmi@isu.edu
11
Thank You!!

## Slide 13
sagarregmi@isu.edu
12
Extra Slides
Plots: Filterwise Average Asymmetry Distributions for 20 filter cycle(7.0 nA LL)
The New ADC Board

## Slide 14: Timing Setup
sagarregmi@isu.edu
13
Timing diagram for PMT non-linearity bench tests. The ADC samples the data only during the tmeas time window

## Slide 15: MOLLER Intro…
sagarregmi@isu.edu
Components
CEBAF beam: ~90% longitudinally polarized, 11 GeV, Helicity reversal @ 960 Hz
Toroidal Spectrometer: Upstream and downstream, focuses the full momentum range of scattered particles ranging 2-8 GeV.
Tracking Chambers
Integrating detector: Main int det, Shower-max detector, Pion background detector
Aux detectors: SAMs, LAMs, Scanners
14

## Slide 16: The weak mixing angle
sagarregmi@isu.edu
15
Fig: The weak mixing angle from the various experiments vs the energy scale. The running blue curve is the standard model prediction. The proposed MOLLER experiment is shown with the expected error bar.

## Slide 17: Spectrometer
sagarregmi@isu.edu
16
Left 2 plots: Relationship between the lab and COM angles for Ebeam=11 GeV
Right plot: Primary acceptance for Moller electrons