## Slide 1: The IFMIF-DONES Accelerator: Target & Irradiation Challenges and Future Research Opportunities
Claudio Torregrosa-Martín
On behalf of the IFMIF-DONES Team

## Slide 2: List of Authors
A. Ibarra, W. Królas, F. Arbeiter, S. Becerril, D. Bernardi, L. Buiglins, J. Castellanos, M. Cappelli, P. Cara, T. Dézsi, A. Diaz, M. García, J. Gutiérrez, D. Jimenez-Rey, M. Luque, J. Maestre, F. Martín-Fuertes, J.C. Marugán, C. de la Morena, G. Miccichè F. S. Nitti, C. Oliver, F. Ogando, I. Podadera, Y. Qiu, D. Regidor, D. Sánchez-Herranz, T. Tadić, M. Vázquez, U. Wiącek, M. Weber

## Slide 3: OUTLINE
Introduction to the IFMIF-DONES Facility
The Irradiation Modules
Summary
The IFMIF-DONES Accelerator
The Lithium Target
Future Research Opportunities
1
2
3
5
4
6

## Slide 4: WHY IFMIF-DONES
One of the main differences between ITER and DEMO is the radiation dose: at DEMO more than one/two orders of magnitude higher
1-3 dpa/lifetime
20-50 dpa/lifetime

## Slide 5: Radiation Damage to Fusion Materials
Radiation makes materials swell
Normalized neutron energy spectrum DEMO first-wall
A fusion-like neutron source is needed!
It makes them brittle!
Radiation damage processes are very dependent on the neutron energy spectrum
Fusion neutrons spectrum differ significantly from the ones of Fission Reactors and Spallation Neutron Sources

## Slide 6: A fusion-like neutron source
We need to produce fusion-like neutrons with:
  Intensity large enough for accelerated testing (as compared to DEMO)
  Damage level above the expected operational lifetime
  Irradiation volume large enough for the characterization of the macroscopic properties of the materials of interest required for the engineering design of DEMO (and the Power Plant)
> 10 dpa(Fe)/fpy
20 dpa(Fe) in 1.5 y
50 dpa(Fe) in 3.5 y
300 cm3
The most feasible approach based on Li(d,xn) sources
The IFMIF project since 90´s
Requirements based on EU DEMO needs
The
IFMIF-DONES Project
Broader Approach IFMIF/EVEDA since 2007 (LIPAc & IFMIF IIEDR)
EUROfusion WPENS since 2015

## Slide 7: What is IFMIF-DONES?
IFMIF-DONES is a combination of:
  High Intensity Deuteron Accelerator
  Liquid Lithium Loop Facility
  Irradiation Modules to house material specimens in a controlled environment
  IFMIF-DONES is an accelerator-based fusion-like neutron source to be used for the qualification of the materials for the DEMO Fusion Reactor

## Slide 8: The IFMIF-DONES site in Granada, Spain
The site is located at Escúzar -18 km southwest from Granada city- Spain

## Slide 9: The Facility
Irradiation Facility delivering up to 1015 n/cm2/s (1017 n/s) with a broad peak at 14 MeV.
Operational at least 255 days/year (70% Operational Availability).
~100 m
LINAC Accelerator of Deuterons
40 MeV
125 mA, CW
5 MW output power
Li Target, Irradiation Modules and Test Cell
Li Loop and Oil Cooling Loops
100 l/s Li
Strict Li Purity Control
Deuteron Beam

## Slide 10: AREAS FOR NON-FUSION EXPERIMENTS
R160
R160: 300 m2 Area
Same floor as the Test Cell
Profiting from a continuous neutron beam from the TC via a Neutron Beam Aperture/Collimator (Ø20-12 cm)
Estimated neutron flux at the duct exit:
2∙10¹⁰ n/cm²/s, 99.5% of fast neutrons, En > 100 keV
R026: 2100 m2 Area
Ground floor, below the Accelerator Vault
Profiting from Pulsed Deuterons beams via a beam extraction from the accelerator
Deuteron extraction at 40 MeV (0.1 % intensity)
Experiment using deuterons & pulsed neutrons (by neutron production with a Target)
High flux Modules
Medium Flux Modules
Low Flux Modules
OTHER FUSION
TEST MODULES

## Slide 11: OUTLINE
Introduction to the IFMIF-DONES Facility
The Irradiation Modules
Summary
The IFMIF-DONES Accelerator
The Lithium Target
Future Research Opportunities
1
2
3
5
4
6

## Slide 12: A 5 MW Power LINAC
Main HEBT @ TIR
HEBT-S1
HEBT-S2
HEBT-S3
HEBT-BDTL
HPBD
CM#5
CM#4
CM#3
CM#2
CM#1
MEBT
RFQ
Injector
Beam footprint at the Target
Highest current LINAC
Top H+ and D+ Injector Performance needed
Longest RFQ
Windowless liquid Li target
87% availability
High residual doses in the area
Very challenging beam Diagnostics and Beam Loss Monitoring
D+ CW beam
125 mA / 40 MeV  5 MW
Towards the Target
LIPAc Acceletaror (5 MeV) built in Rokkasho (Japan) as integrated Prototype of the:
Injector
LEBT
RFQ
Dump
1st Cryomodule

## Slide 13: OUTLINE
Introduction to the IFMIF-DONES Facility
The Irradiation Modules
Summary
The IFMIF-DONES Accelerator
The Lithium Target
Future Research Opportunities
1
2
3
5
4
6

## Slide 14: The Target/Test Cell Environment
Concrete
Shielding
Li Loop Cell
(LLC)
Target Interface Room
(TIR)
Primary Line (beam line)
Secondary Line (diagnostics)
Inlet/Outlet Li Pipes
Test Cell
Target Vacuum Chamber
Irradiation Module
(HFTM/
STUMM)
Li Loop Cell
(LLC)
~3 m
Radiation in the exposed zone of Irradiation Modules:
104 MGy/fpy
30 MGy/fpy (at the connectors bridge)
Radiation in around the Target Vacuum Chamber:
100 MGy/fpy
Radiation in the Target Interface Room (TIR):
1-2 MGy/fpy

## Slide 15: Liquid Lithium Target Challenges (1)
E. Wakai et al., Nucl. Mater. Energy 9 (2016), 278-285
Liquid Li jet flowing at 15 m/s,
thickness = 25+/- 1 mm.
Jet stability & Target position is critical to avoid Back-Plate rupture.
Li Target: Jet flowing at 15 m/s, thickness = 25+/- 1 mm.
Jet stability & Target position is critical to avoid Back-Plate rupture.
Target Vacuum Chamber shall be within 10-4 -10-5 mbar.
Full Remote Handling replacement/maintenance.
Time for 75% BP material melting
Time for BP vaporization onset
Time for BP melting onset
Time for 25% BP material melting
A Li thickness reduction
> 4/5 mm would lead to power deposition in the Back-Plate
Failure/melting could take place in less than 10 ms

## Slide 16: Technologies to monitor the Li Target thickness
Laser measurement: Amplitude-modulated Continuous-wave Light Detection and Ranging (AMCW LiDAR)
High-frequency modulation at 550 MHz
100 microns accuracy
Acquisition times within 1 ms per point and <100 ms per line.
Compatible with the 1-2 MGy/fpy in the TIR
Radiofrequency (mmWave radar techniques)
Linear Frequency Modulated Continuous Wave (LFM-CW) Radar

## Slide 17: Liquid Lithium Target Challenges (2)
Monitoring the vacuum pressure in the TVC
Avoid Li boiling !! (above vapor saturation pressure)
Managed by Ar injection system to achieve a differential vacuum pressure in the accelerator
Very high radiation (100 MGy/fpy) disables the installation of vacuum gauges in the TVC
Beam-on-Target Position
Beam foot-print Area (20 x 5 cm)
(Up to 25 cm wide counting with the tales)
Only 5 mm margin in each side to avoid touching the wall
Y. Momozaki et al. Journal of Radioanalytical and Nuclear Chemistry volume 305, pages 843–849 (2015)
Keeping vacuum in the Target chamber within 10-4 – 10-5 mbar
Camera pointing to the jet through an optical path from the TIR diagnostic port.
Recording interaction between the beam and the lithium surface (OTR) and residual gas fluoresce

## Slide 18: Liquid Lithium Target Challenges (3)
Yearly replacements of the Target is required
It must be fully done by remote tele-manipulators and cranes.
  x2 large vacuum flange connections
  Inlet/outlet Li pipes
  Alignment of the assembly

## Slide 19: The Liquid Lithium Loop and 2nd/3rd Oil Loops
100 l/s of Li at 300 °C.
EMP Pump, heaters, Heat-Exchangers and Dump tank.
Li Loop room in Inert atmosphere (Ar)
Impurity Control Loop
Extremely low impurities requires
Purification loops with “Traps” Cold Traps, Tritium/H traps, N Traps
Permanent Magnets Electro-Magnetic Pump
ELTL Lithium Loop Prototype in Oarai (Japan)
Liquid Li loop at 40 l/s
1:1 for Loop height and jet velocity (15 m/s)
Operating within 2010-2014
Fluid-Dynamic stability of the Li jet validated (1560 h)
E. Wakai et al., Nucl. Mater. Energy 9 (2016), 278-285

## Slide 20: OUTLINE
Introduction to the IFMIF-DONES Facility
The Irradiation Modules
Summary
The IFMIF-DONES Accelerator
The Lithium Target
Future Research Opportunities
1
2
3
5
4
6

## Slide 21: Challenges of the Irradiation Modules (1)
Characterizing the radiation field with high spatial and time resolution
5·1014/1015 n/cm²s
Up to 14 MeV (80% above 1 MeV). Fast neutrons dominating
Spatial resolution: 10 mm
Time resolution: 10 µs.
Maintain n-field calibration during 1 year
104 MGy/fpy exposed zone, 30 MGy/fpy at connectors plate
Surface of 50x40 cm for integrating remote handling connectors for more than 330 different signals.
Very low current signals, long cables, EMC sensibility..
Online/Offline Radiation Diagnostics
SPNDs (Self-powered Neutron Detectors). (compatible with high temperatures)
µFission (U238/U235) & Ionization chambers
Activation foils
Pneumatic Rabbit transporting activation balls (for the STUMM)

## Slide 22: Challenges of the Irradiation Modules (2)
Irradiation Capsules with miniaturized materials Specimens
32 irradiation capsules, 850 specimens can be irradiated to 12 – 25 dpa/fpy.
Irradiation temperatures shall be controlled within 250 – 550 ºC. -> Heaters and thermocouples for each capsule.
Filled by Liquid Na, homogeneous Temperature
Necessary to know the integral radiation to which the specimens have been exposed. (SPNDs and activation foils)
Prototyping Activities of Irradiation Modules at KIT

## Slide 23: OUTLINE
Introduction to the IFMIF-DONES Facility
The Irradiation Modules
Summary
The IFMIF-DONES Accelerator
The Lithium Target
Future Research Opportunities
1
2
3
5
4
6

## Slide 24: IFMIF-DONES Users Community!
A Users Community is being launched to enhance Experiment Proposals for both fusion and non-fusion applications in multidisciplinary fields
1st IFMIF-DONES Users Workshop in 2022 -> Online
2nd IFMIF-DONES Users Workshop in 2023 -> Granada
3rd IFMIF-DONES Users Workshop in 2024 -> to take place in Zagreb, Croatia, 1st-2nd October 2024
Some examples of proposals/applications in the next slides
Each year, a “DONES Users Workshop“ is organized
More info in: https://ifmif-dones.es/dones-users-2/

## Slide 25: Additional Irradiation Modules in the Test Cell
1) Experiments Irradiating Material Specimens
Creep Fatigue Test Module (CFTM)
2) Experiments Irradiating Components
Activated Corrosion Product Test Module (ACPTM)
Irradiation of Instruments and Diagnostics
In-situ Ceramic Breeder Test Modules
3) Experiments on Tritium Breeders Materials
Ceramic breeders (Li4SiO4)
Liquid Metal breeders (like PbLi)
4) Radioisotopes Production

## Slide 26: Complementary Experiments using neutrons in the R160
2) Fundamental particle physics studies with cold and ultra-cold neutrons.
3) Neutron scattering experiments with moderated beams
R160
Continuous neutron beam 2∙10¹⁰ n/cm²/s, 99.5% of fast neutrons, En > 100 keV
1) Experiments with samples exposed to neutrons:
Activation of small samples by neutrons
Characterization of materials by radiation analysis
Neutron imaging
Fast neutron irradiation of components, devices or bio-samples

## Slide 27: Complementary Experiments using neutrons in the R026
2) Experiments using Pulsed Deuterons Beams
  Production of radionuclides with intense deuteron beams
  Research on Deuteron-induced reactions
Experiments using Pulsed Neutron Beams (nTOF):
(via deuteron/neutron converter)
  Neutron Cross-section measurements
  Gamma-spectroscopy of the nuclei produced in fast-neutron-induced fission reaction
  Half-life measurements of long-lived isotopes
Profiting from Pulsed Deuterons beams via a beam extraction from the accelerator at 40 MeV (0.1 % total intensity -> 0.125 mA)

## Slide 28: OUTLINE
Introduction to the IFMIF-DONES Facility
The Irradiation Modules
Summary
The IFMIF-DONES Accelerator
The Lithium Target
Future Research Opportunities
1
2
3
5
4
6

## Slide 29: SUMMARY
IFMIF-DONES will be a 1st kind and 1st Class Accelerator Facility.
Its technological challenges are pushing the state-of-the-art in many aspects of accelerator, targets and instrumentation.
Extensive R&D programs and ambitious prototypes have been carried out and are ongoing to develop the technical solutions required (and many more will follow).
Aiming at reinforcing links and collaborations between Accelerators, Fusion and Fission Communities
Pursuing a wide range of Applications in an extensive variety of multidisciplinary nuclear fields, in addition to fusion.

## Slide 30
Thanks very much!!
Questions?