## Slide 1: Characterizing the AFP Spin Flipper for the Nab Experiment
Rebecca Godri
27 September 2024

## Slide 2: Outline
Introduction to the SNS and the Nab experiment
Importance of polarimetry within the Nab experiment
Nab spin flipper
  Design
  Adiabaticity
Characterization effort
  Devices
  Results
Future work

## Slide 3: The Spallation Neutron Source
Uses a pulsed beam of negative hydrogen ions
Beam is accelerated and circulates around an accumulator ring approximately 1200 times; this forms the proton pulse
Proton pulse reaches steel target filled with liquid mercury every 60 Hz
Protons hit the nucleus of the mercury atom and release neutrons which are guided to various beamlines

## Slide 6: The Nab experiment
Neutron beta-decay experiment (Neutron “a” and “b”)
  a: electron-neutrino correlation parameter
  b: Fierz interference term
Goal: to make a precise determination of
  a () and b ()

## Slide 7: Experimental highlights
Proton momentum is determined from TOF in magnetic field
Plotting the proton phase space creates a teardrop shape

## Slide 9: Effect of Beam Polarization within the Nab Experiment
⇒
If σn≤ 2 x 10-5
= ☺
If σn> 2 x 10-5
To meet our experiment goal:
Must measure beam polarization and if found to be non-zero, it must be periodically reversed during data taking

## Slide 10: The Nab Spin Flipper
The Nab spin flipper is an Adiabatic Fast Passage (AFP) spin flipper
Neutrons align themselves with the gradient field and precess
Spin flip occurs when the Larmor frequency of the neutron is equal to the frequency of the oscillating field
13B is a polychromatic beam; we want all wavelengths to flip

## Slide 11: The Nab Spin Flipper
The Nab spin flipper is an Adiabatic Fast Passage (AFP) spin flipper
Neutrons align themselves with the gradient field and precess
Spin flip occurs when the Larmor frequency of the neutron is equal to the frequency of the oscillating field
13B is a polychromatic beam; we want all wavelengths to flip

## Slide 12: Characterizing the spin flipper at FNPB
Second characterization effort
Monochromatic beam (4.45 Å and 8.9 Å)

## Slide 13
Task: perform a series of transmission measurements with different components on/off, solve for individual device efficiencies

## Slide 15: Devices used in characterization measurements
Neutron filtering devices
Collimators
Nab spin flipper
Mezei spin flipper
8-pack detector

## Slide 16: Neutron filtering devices
Supermirror polarizer
  Alternating layers of ferromagnetic and nonmagnetic materials
    Magnetic: large nuclear and magnetic scattering lengths
    Non-magnetic: scattering length close to zero
    These layers create a highly reflective polarizing mirror for one spin state
  One spin state is reflected (absorbed) while the other is refracted (transmitted)
S-bender
  Essentially two C-shaped supermirror polarizers
  Serves as our analyzing device

## Slide 17: Mezei spin flipper
3 main components
Inner coil – spin flip coil which provides a perpendicular guide field
Outer frame – provides a continuous guide field
Outer coil – cancels the guide field so the neutron effectively precesses about the perpendicular guide field from the inner coil
Notes: Explain why we have a second spin flipper

## Slide 18: 8-pack detector
Eight 3He filled tubes oriented horizontally
  Neutrons capture on 3He
A neutron event is determined via resistive anode charge division
Preamplifiers convert signals into voltages to determine the neutron interaction location

## Slide 19: Four measurement sequence
N0: no flips
N1: Nab SF flips
N2: Mezei SF flips
N12: Both flippers flip
Two configurations (upstream and downstream)
Two wavelengths (4.45 Å and 8.9 Å)
High field
Low field

## Slide 20: Choppers
Rotating disks with a pizza slice-shaped opening
TOF delay “chops” neutrons with an undesired wavelength
4.45 Å
4.45 Å
8.9 Å
8.9 Å
4.45 Å

## Slide 21: Upstream (4.45 Å)
Both off | Nab on | Mezei on | Both on | f (Nab) | f’ (Mezei) | Combined efficiency
53775.65 ±231.90 | 8186.298 ±90.48 | 4103.63 ±64.06 | 49750.14 ±223.05 | 0.918 ±0.0031 | 1.001 ±0.0037 | 0.857 ±0.0063
54988.64 ±234.50 | 8215.24 ±90.64 | 3906.08 ±62.50 | 50088.73 ±223.81 | 0.910 ±0.0031 | 0.994 ±0.0036 | 0.878 ±0.0064

## Slide 22: Downstream (4.45 Å)
Both off | Nab on | Mezei on | Both on | f (Nab) | f’ (Mezei) | Combined efficiency
53038.38 ±230.30 | 27542.6 ±165.96 | 4386.81 ±66.23 | 25778.61 ±160.56 | 0.482 ±0.0024 | 0.9195 ±0.0058 | 0.995 ±0.0127
51661.98 ±227.29 | 29121.28 ±170.65 | 4534.72 ±67.34 | 26838.33 ±163.82 | 0.476 ±0.0025 | 0.995 ±0.0074 | 0.847 ±0.0118

## Slide 23: Investigating non-uniformity
 | Both off | Nab on | Mezei on | Both on | f (Nab) | f’ (Mezei) | Combined efficiency
Left | 57587.36 ±239.97 | 6109.36 ±78.16 | 4087.98 ±63.94 | 54059.16 ±232.51 | 0.948 ±0.0031 | 0.985 ±0.0034 | 0.892 ±0.0061
Right | 57885.96 ±240.595 | 7200.81 ±84.56 | 4306.53 ±65.62 | 54361.07 ±233.15 | 0.940 ±0.0031 | 0.994 ±0.0034 | 0.872 ±0.0060

## Slide 24: Investigating non-uniformity
 | Both off | Nab on | Mezei on | Both on | f (Nab) | f’ (Mezei) | Combined efficiency
Left | 57587.36 ±239.97 | 6109.36 ±78.16 | 4087.98 ±63.94 | 54059.16 ±232.51 | 0.948 ±0.0031 | 0.985 ±0.0034 | 0.892 ±0.0061
Right | 57885.96 ±240.595 | 7200.81 ±84.56 | 4306.53 ±65.62 | 54361.07 ±233.15 | 0.940 ±0.0031 | 0.994 ±0.0034 | 0.872 ±0.0060
Center | 54988.64±234.497 | 8215.24±90.64 | 3906.08±62.499 | 50088.73±223.81 | 0.9099 ±0.0031 | 0.994 ±0.0036 | 0.878 ±0.0064

## Slide 25: Future work
Further data analysis
Further characterization at HFIR (late 2024)
Characterization and implementation on BL-13B (2025)
  Optimize S-bender parameters
  Design and build an in-situ 3He analyzer system
  Optimize new 3He tube characteristics for final monitor
This material is based upon work supported by the U.S. Department of Energy, Office of Science, Office of Nuclear Physics Fundamental Symmetries program under Award Number DE-FG02-03ER41258
P. Jiang (ORNL)

## Slide 26
The Nab collaboration
Main project funding:
Nab collaborating institutions:
University of South Carolina, Universität Karlsruhe (TH), Universidad Nacional Autónoma de México, Western Kentucky University

## Slide 27: Backup slides
:)

## Slide 28: Tuning polarizing devices

## Slide 29: Mezei flipper settings
S-bender installed
Peaks indicate non-flip states
Valleys indicate flipped states

## Slide 30: Upstream (8.9 Å)
 | Both off | Nab on | Mezei on | Both on | f_Nab | f_Mezei | Combined efficiency
8/23/24 Run 1 | 91274.74 ±302.12 | 21037.28 ±145.04 | 20076.65 ±141.69 | 87176.68 ±295.26 | 0.964 ±0.0032 | 0.978 ±0.0033 | 0.664 ±0.0044
8/23/24 Run 2 | 85888.46 ±293.07 | 21886.78 ±147.94 | 20500.76 ±143.18 | 86472.31 ±294.06 | 0.994 ±0.0035 | 1.02 ±0.0037 | 0.5997 ±0.0042
8/29/24 Run 1 | 79093.33 ±281.24 | 33256.55 ±182.36 | 19828.02 ±140.81 | 89460.47 ±299.099 | 0.974 ±0.0049 | 1.2596 ±0.0066 | 0.423 ±0.0037
8/29/24 Run 2 | 78125.93 ±279.51 | 31500.45 ±177.48 | 19742.69 ±140.51 | 90021.55 ±300.04 | 1.001 ±0.0040 | 1.254 ±0.0064 | 0.425 ±0.0037

## Slide 31: Upstream (8.9 Å)

## Slide 32: Upstream (4.45 Å and 8.9 Å)

## Slide 33: Downstream (8.9 Å)
Both off | Nab on | Mezei on | Both on | f (Nab) | f’ (Mezei) | Combined efficiency
78944.57 ±280.97 | 45366.17 ±212.99 | 19642.75 ±140.15 | 62986.3 ±250.97 | 0.649 ±0.0029 | 1.145 ±0.0080 | 0.488 ±0.0055
80225.0 ±283.24 | 46865.66 ±216.48 | 20147.7 ±141.94 | 58977.54 ±242.86 | 0.601 ±0.0028 | 1.082 ±0.0075 | 0.529 ±0.0061

## Slide 34: Downstream (4.45 Å and 8.9 Å)