## Slide 1: Amplifier Additive1 Phase Modulation Noise Characterization and Testing and Theoretical Discussion of White and Flicker Noise Processes in both Small and Large Signal Regions of Operation* Mini-Circuits
Let’s address these common questions:
What is it?
How is it measured?
How is it specified?
Test Set Description
Additive Phase Noise Theory
Small Signal Behavior
Large Signal Behavior
AM-to-PM Distortion Influence
Broadband and Cascade Design Tradeoffs
Joseph L. Merenda
V.P. of Engineering
JoeM@minicircuits.com
*Class A, AB Only. 1:flicker and white
Low-Level Radio Frequency (LLRF) Workshop 2025
Id#107

## Slide 2
R&S FSWP:
What’s Inside?
All images courtesy of Rohde & Schwarz GmbH & Co. KG
Cancel un-correlated noise not common to Channels 1 and 2.
RE: AM Noise, IM: PM Noise

## Slide 3
Shielded DC Cables (3)
EMI Enclosure
FSWP Analyzer
Short RF I/O Cables
Vibration isolated table
Class 100k clean room
Filtered AC power
Test Lab

## Slide 4
For best results:
APM/AAM measurements must be performed with short, well matched, low loss, shielded and phase stable RF cables and well shielded DC cables.
Care should be taken when torquing all cable connectors to ensure they are tight and that any mechanical (torsional) stress in the cables is eliminated.
The DUT (and in certain instances the test equipment itself) must be placed inside a suitable EMI shielded enclosure.
It is also very important that the DUT and test equipment be well isolated from any external, mechanical vibration and AC line noise.
The ambient, DUT and test equipment temperatures be stable and held constant during the measurement.
Be sure not to disturb or touch anything during the measurement process!
APM: Additive Phase Modulation Noise
AAM: Additive Amplitude Modulation Noise

## Slide 5
Signal Source Output
(To DUT)
Fc, Po
Input Step Attenuator
(To DUT)
Input Fixed Pad
(To DUT)
RF Input
(From DUT)
Output Step Attenuator
(From DUT)
Output Fixed Pad
(From DUT)
Short RF I/O Cables
Test Lab

## Slide 6
For our analyzer:
Fixing the signal source power at the optimum level and using an external, mechanical step attenuator and fixed pad on the signal source output port to set the power level into the DUT, provides more consistent APM performance compared to that achieved using the analyzer’s internal attenuator.
We use a second mechanical step attenuator and fixed pad after the DUT, on the input to the analyzer’s RF input port and adjust the attenuation to maintain a constant signal power into the analyzer’s RF input port.
This method provides for enhanced accuracy and repeatability and ensures that maximum input power levels are never exceeded.
Ideally, the RF input port signal power is adjusted and maintained at a fixed power level, between ≈+7 to +13 dBm for all input attenuator steps.

## Slide 7
Measurement Parameters include the following:
• Carrier Frequency: As required (0.1-8 GHz, typical).
• Signal Source Output power level: +13 dBm, or max allowable.
• Measurement: Additive Noise.
• Offset Frequency: 100 Hz to 1 MHz (as required).
• Spot Noise: 100, 1,000, 10,000, 100,000, 1,000,000 Hz.
• Resolution Bandwidth: 1%.
• Cross-Correlation Factor: 100 to 10,000 (as required).
• Cross-Correlation Threshold: 15 dB.
• Spurious Threshold: 6 dB.
• Sweep Average Count: 0.
• Smoothing Value: 1%.
• RF Input: Fixed between +7 dBm to +13 dBm.
Key Settings

## Slide 8
PSD of Random Phase ф(t) = Sф (f) = b0 + b-1/f [rad^2/Hz] = FkTo/Pin + b-1/f [rad^2/Hz]
and L(f) [rad^2/Hz] = ½ Sф (f) [rad^2/Hz]; L(f) [dBc/Hz] = 10*Log10(Sф (f)) - 3dB
Single Device Small Signal Operation
(L(f)= ½ Sф (f) for delta ф_peak <~0.2 rad)
[2]
What is Amplifier Additive Phase Noise?
Flicker(parametric) White(additive)

## Slide 9
Amplifier additive 1/f and White noise degrades oscillator noise, Jitter and cascade S/N ratio
Converting Oscillator Phase Noise to Time Jitter
by Walt Kester
[2]
Why do we care?
Google search
“Time fluctuations are critical in the RF cavities of particle accelerators where they cause intensity loss or energy loss in the beam.” F. Tecker, H. Damereau.
BER

## Slide 10
MDS: Minimum Discernable Signal at Input
MDS = FkT0B. B=1Hz, T0=290K, k=1.38x10-23W*S/K
b0 = FkT0/P0=MDS/P0 [rad^2/Hz]
Pout
MDSout: Minimum Discernable Signal at Output
MDSout = MDS*G = FkT0G, B=1 Hz
b0_out = MDSout /Pout = FkT0G/Pout => FkT0G/Psat
b0_out = b0 , Pout = G P0
Cascade: Ftot=F1+(F2-1)/G1+(F3-1)/G1G2+--- => maximize G1
G, F
Optimum APM:
  Minimize Gain and Noise Figure
  Maximize Pout (Psat)
Good Luck!!
Cascade!!
What is White Noise?
[2]
Typ. Rule of Thumb
(does it hold?)

## Slide 11
What is Flicker or 1/f Noise?
Flicker noise with 1/f power law (~Size of Event/Number of Events) observed in almost all electronic materials
Also observed in mechanical, biological, geological, economic and musical systems
How is it generated? Several theories exist.
Fluctuations of charge carriers due to trapping in surface states
Fluctuations in conductivity due to fluctuations in bulk mobility*
Amplifiers: Dependent upon DC current and voltage, device type, process node size, power density,
fluctuations in I/O impedance and insertion gain and phase,
I/O Impedance matching criteria, input power and carrier frequency.
Also see Dr. Ali Hajimiri
For HEMT Devices*:
ns(T): electron gas density
N: number of electrons
(by Hooge)
For HBT Devices:
High N+*W*L- and
Low I_EQ
High R_E (un-bypassed)
High Vcb
Low Cbe, Cce
Low G_dc
DC Filtering
Passives
HBT “always” better than HEMT?
“Flicker noise in AlGaAs/GaAs of high electron
mobility heterostructure field-effect transistor at
cryogenic temperature” [11]

## Slide 12
What is Flicker or 1/f Noise?
HBT “always” better than HEMT?
Trap density increases as Lg decreases which increases 1/f noise

## Slide 13
What is Flicker or 1/f Noise?
Amplifiers: Dependent upon DC current and voltage, device type, process node size and power density
Fluctuations in impedance and insertion gain and phase. “1/f Noise in Proton Irradiated SiGe HBTs” [10]
Ae=We*Le*N
0.5x(0.50x2)
0.5x(1.25x2)
0.5x(5.00x2)

## Slide 14
HBT “always” better than HEMT
dashed: pHEMT
Solid: HBT
For pHEMTs:
~Increasing Lg
0.15um<Lg<0.5um
(45GHz>Fo>50MHz)
Lg=0.25um
Lg=0.15um
Lg=0.5um
For HBTs:
We=2um
(27GHz>Fo>DC)
HBT “always” better than HEMT?
Trap density increases as Lg decreases which increases 1/f noise

## Slide 15
LVA-273PN+ APM Performance at 2, 3, 4, 5, 6, and 7 GHz Carrier Frequencies. Pin=-7dBm.
(BLK:2GHz, RED:3GHz, GRAY:4GHz, ORANGE:5GHz, GREEN:6GHz, BLUE:7GHz)
L(f) [dBc/Hz] = -177 [dBm/Hz] +NF [dB] – Pin [dBm]
(valid for small signal operation)
Noise is Additive
Non-DUT related, will discuss next.
How “Low” is Low Power?
(HBT)
Let’s make some measurements!

## Slide 16
Test Set through path APM performance at 2, 3, 4, 5, 6, and 7 GHz carrier frequencies
(BLK:2GHz, RED:3GHz, GRAY:4GHz, ORANGE:5GHz, GREEN:6GHz, BLUE:7GHz)
Cross-Spectral Collapse?
Phase Length
(de-correlation)
Non-DUT related. Adjust source power.
Increasing F_Carrier
Non-DUT related.
How “High” is High Power?
Let’s make some measurements!
-177dBm/Hz-13dBm=-190dBc/Hz
Use short RF cables.

## Slide 17
Examples of Cross-Spectral Collapse
Cross Correlation: Cancel un-correlated noise not common to Channels 1 and 2.
Anti-Correlated AM noise
AM/PM Noise Cancellation
Anti-Correlated Supply noise

## Slide 18
LVA-273PN+ small signal APM vs. offset frequency at various input powers from:
P_Opt – 4 dB (gray-top trace) to P_Opt (blue-bottom trace), 1dB steps.
Increasing Pin (SS)
1/f Insensitive to Pin (SS)
Small Signal Operation
L(f) [dBc/Hz] = -177 [dBm/Hz] +NF [dB] – Pin [dBm]
(valid for small signal operation)
Noise is Additive
(HBT)

## Slide 19
GVA-84+ Flicker-to-white noise cutoff offset frequency shift with increasing input powers. From: Pin_Opt (Trace SN4-bottom) to Pin_Opt + 4 dB (Trace M1-top), 1 dB steps.
Increasing Pin (LS)
1/f Sensitive to Pin (LS)
Large Signal Operation
L(f) [dBc/Hz] = -177 [dBm/Hz] +NF(Pin) [dB] – Pin [dBm]
(approximate for large signal. Need AM and PM)
Noise is Parametric
(HBT)
??

## Slide 20
At 2 GHz Carrier
L(f) [dBc/Hz]* = -177 [dBm/Hz] +NF(Pin) [dB] – Pin [dBm] [4]
(approximate for large signal.)
Additive
Parametric
Additive
Parametric
Key Equation!
*White noise region

## Slide 21
LVA-273PN+: 2 GHz Additive Phase Noise (APM) [dBc/Hz] at 10 kHz and 1 MHz offsets, versus Input Power [dBm]. (Pin_Opt = ~+11 dBm, Pin_1dB = ~+1 dBm).
GVA-84+: 2 GHz Additive Phase Noise (APM) [dBc/Hz] at 10 kHz and 1 MHz offsets, versus Input Power [dBm]. (Pin_Opt = ~+1 dBm, Pin_1dB = ~+3 dBm).
Mini-Circuits Proprietary
Mini-Circuits Proprietary
G = 18dB
NF = 6 dB
P1dB = +21 dBm
OIP3 = +37 dBm
Darlington-Cascode
G = 17dB
NF = 8 dB
P1dB = +17 dBm
OIP3 = +27 dBm
Distributed*
2 um InGaP/GaAs HBT
2 um InGaP/GaAs HBT
*Hint: think parallel combining!

## Slide 22
GVA-62+: 2 GHz Additive Phase Noise (APM) [dBc/Hz] at 10 kHz and 1 MHz offsets, versus Input Power [dBm]. (Pin_Opt = ~+5 dBm, Pin_1dB = ~+6 dBm).
GVA-84+: 2 GHz Additive Phase Noise (APM) [dBc/Hz] at 10 kHz and 1 MHz offsets, versus Input Power [dBm]. (Pin_Opt = ~+1 dBm, Pin_1dB = ~+4 dBm).
Mini-Circuits Proprietary
Mini-Circuits Proprietary
G = 15 dB
NF = 5 dB
P1dB = +20 dBm
OIP3 = +35 dBm
Darlington w/ Re feedback
G = 18dB
NF = 6 dB
P1dB = +21 dBm
OIP3 = +37 dBm
Darlington Cascode
2 um InGaP/GaAs HBT
2 um InGaP/GaAs HBT

## Slide 23
Effect of R_E on APM
GVA-84+:
Idq=110 mA
Gdc=24 dB
P1dB=+21 dBm
NF=6 dB
2 um PWR InGaP/GaAs HBT
Darlington Cascode
R_E84= ???
GVA-62+:
Idq=82 mA
Gdc=15 dB
P1dB=+20 dBm
NF=5 dB
2 um PWR InGaP/GaAs HBT
Darlington w/Re feedback
R_E62= ???(> R_E84)
Mini-Circuits Proprietary
Mini-Circuits Proprietary
Fc=2 GHz
Fc=2 GHz

## Slide 24
AM-to-PM Distortion
(Fc=2GHz)

## Slide 25
GVA-84+:
G = 18dB
NF = 6 dB
P1dB = +21 dBm
OIP3 = +37 dBm
Po_+1dBm=+19dBm
Darlington-Cascode
LVA-273PN+:
G = 17dB
NF = 8 dB
P1dB = +17 dBm
OIP3 = +27 dBm
Po_+10dBm=+22dBm
Distributed
AM-to-PM Distortion
and
APM
(Fc=2GHz)
GOOD
BEST
BETTER
GVA-62+:
G = 15 dB
NF = 5 dB
P1dB = +20 dBm
OIP3 = +35 dBm
Po_+5dBm=+19.5dBm
Darlington w/Re feedback
Studies on small and large signal noise in Solid-State Amplifiers [7]

## Slide 26
GVA-84+:
G = 18dB
NF = 6 dB
P1dB = +21 dBm
OIP3 = +37 dBm
Po_+1dBm=+19dBm
Darlington-Cascode
LVA-273PN+:
G = 17dB
NF = 8 dB
P1dB = +17 dBm
OIP3 = +27 dBm
Po_+10dBm=+22dBm
Distributed
-5deg/dB
-3.6deg/dB
-1.1deg/dB
2GHz
4GHz
6GHz
-0.5deg/dB
-1deg/dB
-1.6deg/dB
2GHz
4GHz
6GHz
-5deg/dB
-3.4deg/dB
-1.3deg/dB
What happens with carrier frequency?
GVA-62+:
G = 15 dB
NF = 5 dB
P1dB = +20 dBm
OIP3 = +35 dBm
Po_+5dBm=+19.5dBm
Darlington w/Re feedback

## Slide 27
AM-to-PM Distortion
G_Compression
At different carrier frequencies
LVA-273PN+:
What happens with carrier frequency?

## Slide 28
L(f) [dBc/Hz] = -177 [dBm/Hz] +NF [dB] – Pin [dBm]
=>AM-to-PM Distortion
10 kHz Offset
100 kHz Offset

## Slide 29
AM-to-PM Distortion
AM-to-PM Distortion
Device Characterization Required
Device Characterization Required
Mini-Circuits Proprietary
Mini-Circuits Proprietary
G = 17 dB
NF = 5 dB
P1dB = +17 dBm
OIP3 = +27 dBm
G = 19 dB
NF = 5 dB
P1dB = +20 dBm
OIP3 = +30 dBm
2 um InGaP/GaAs HBT
2 um InGaP/GaAs HBT

## Slide 30
Mini-Circuits Proprietary
AM-to-PM Distortion
Temperature Characterization Required
AM-to-PM Distortion

## Slide 31
Characterization over Operating Voltage Range may offer improvement in AM-to-PM and APN performance
Additive Phase Noise | 
YELLOW: +6V | 
BLUE: +7V | 
GREEN: +8V | 
ORANGE: +9V | 
ALL Icc = 150mA@+5dBm | 
Pin = +5 dBm | 
Is there anything we can do?

## Slide 32
Characterization over Operating Voltage Range may offer improvement in AM-to-PM and APN performance
Additive Phase Noise | 
YELLOW: +6V | 
BLUE: +7V | 
GREEN: +8V | 
ORANGE: +9V | 
ALL Icc = 150mA@+10dBm | 
Pin = +10 dBm | 
Is there anything we can do?

## Slide 33: Amplifier Additive Phase Modulation Noise Characterization and Testing and Theoretical Discussion of White and Flicker Noise Processes in both Small and Large Signal Regions of Operation* Mini-Circuits
Part 1: What is it? How is it measured? How is it specified?
Part 2: Cascaded LO Drive Amplifiers
2-Series
4-Parallel
Part 3: Cascaded LO Drive Amplifiers
1-Series -> 4-Parallel
AM-to-PM Distortion Influence
Broadband and Cascade Design Tradeoffs
Low-Level Radio Frequency (LLRF) Workshop 2025
*Class A, AB Only
Id#107

## Slide 34
PSD of Random Phase ф(t) = Sф (f) = b0 + b-1/f [rad^2/Hz] = FkTo/Pin + b-1/f [rad^2/Hz]
and L(f) [rad^2/Hz] = ½ Sф (f) [rad^2/Hz]; L(f) [dBc/Hz] = 10*Log10(Sф (f)) - 3dB
[2]
Cascaded Devices Small Signal Operation
+/-3dB*LOG2(m)
*Hint: think Distributed!

## Slide 35
Series Case 1: 6GHz
Amp1: 0 – 2dB Comp
Amp2: 5 – 13dB Comp
ZX60-223LPN+
(LVA-273PN+)
0 dB PAD
2G: NF=8.1,Gs=34dB
6G: NF=4.1,Gs=34dB
SS
SAT

## Slide 36
Series Case 2: 6GHz
Amp1: 0 – 2dB Comp
Amp2: 2 – 7dB Comp
ZX60-223LPN+
6 dB PAD
SS
LS
2G: NF=8.4,Gs=28dB
6G: NF=4.3,Gs=28dB

## Slide 37
Series
Case: 6GHz
Series
Case: 2GHz
X1-6dB-X1 ZX60-233LPN+
X1-6dB-X1 ZX60-233LPN+
X1-0dB-X1 ZX60-233LPN+
X1-0dB-X1 ZX60-233LPN+
SS
SS
LS
LS
SS
SS
SAT
SAT
Green: -0.89/+9.7dBm
Red: +5.5/+15dBm
Red: +5.5/+20.7dBm
Green: -0.89/+15.8dBm
Slope Eq.
Pout=+24.4dBm
Pout=+22.8dBm
Pout=+25.1dBm
Pout=+24.5dBm
tj=10.0 fS,p-p
tj=4.5 fS,p-p
tj=9.8 fS,p-p
tj=4.0 fS,p-p

## Slide 38
Series
Case: 6GHz
Series
Case: 2GHz
X1 ZX60-233LPN+
X1 ZX60-233LPN+
SS
SS
X1-6dB-X1 ZX60-233LPN+
X1-6dB-X1 ZX60-233LPN+
SS
SS
LS
LS
Green: -0.89/+9.7dBm
Red: +5.5/+15dBm
Red: +5.5dBm
Green: -0.89dBm
Pout=+24.5dBm
Pout=+22.8dBm
Pout=+15.8dBm
Pout=+20.7dBm
tj=7.8 fS,p-p
tj=9.8 fS,p-p
tj=3.3 fS,p-p
tj=4.0 fS,p-p

## Slide 39
ZN4PD-02183-S+
X4 ZX60-223LPN+
ZN4PD-02183-S+
Parallel Case 1: 6GHz
X4 Amp1: 1 – 3dB Comp
LS

## Slide 40
Parallel
Case: 6GHz
Parallel
Case: 2GHz
X1 ZX60-233LPN+
X4 ZX60-233LPN+
X1 ZX60-233LPN+
X4 ZX60-233LPN+
SS
SS
LS
LS
Gray: +3.1dBm
Red: +4.9dBm
Red: +3.8dBm
Noise Floor Limited
Noise Floor Limited
Red: +5.5dBm
Pout=+20.7dBm
Pout=+25.0dBm
Pout=+26.2dBm
Pout=+19.0dBm
tj=7.8 fS,p-p
tj=3.3 fS,p-p

## Slide 41: Amplifier Additive Phase Modulation Noise Characterization and Testing and Theoretical Discussion of White and Flicker Noise Processes in both Small and Large Signal Regions of Operation* Mini-Circuits
Part 1: What is it? How is it measured? How is it specified?
Part 2: Cascaded LO Drive Amplifiers
2-Series
4-Parallel
Part 3: Cascaded LO Drive Amplifiers
1-Series -> 4-Parallel
AM-to-PM Distortion Influence
Broadband and Cascade Design Tradeoffs
Low-Level Radio Frequency (LLRF) Workshop 2025
*Class A, AB Only
Id#107

## Slide 42
Series/Parallel Case 1: 6GHz
X1 Amp1 (A1) – X4 Amp1 (A2)
X1:1 – 5dB Comp. X4(each): 5 – 10dB Comp
X1 ZX60-223LPN+
0 dB PAD
ZN4PD-02183-S+
X4 ZX60-223LPN+
ZN4PD-02183-S+
SAT
LS
2G: NF=8.1,Gs=33dB
6G: NF=4.1,Gs=33dB

## Slide 43
Series/Parallel Case 2: 6GHz
X1 Amp1 – X4 Amp1
X1:1 – 2dB Comp. X4(each): 1 – 2dB Comp
X1 ZX60-223LPN+
8 dB PAD
ZN4PD-02183-S+
X4 ZX60-223LPN+
ZN4PD-02183-S+
SS
SS
2G: NF=8.5,Gs=25dB
6G: NF=4.6,Gs=25dB

## Slide 44
X1 - 0dB - X4 ZX60-233LPN+
(A1: large signal. A2: SAT)
X1 - 8dB – X4 ZX60-233LPN+
(A1:small signal
A2: small signal)
6GHz
X1 ZX60-223LPN+
8 dB PAD
ZN4PD-02183-S+
X4 ZX60-223LPN+
ZN4PD-02183-S+
“Best” Pin
SAT
LS
SS
SS
X1 ZX60-223LPN+
0 dB PAD
ZN4PD-02183-S+
X4 ZX60-223LPN+
ZN4PD-02183-S+
Red: +10.6/+16.8dBm
Blue: +0.4/+10.2dBm
Red: +4.5/+5.8dBm
6G: NFs=4.1,Gs=33dB
6G: NFL=10.4,GL=28dB
6G: NFs=4.6,Gs=25dB
6G: NFL=11.5,GL=22dB
6G: NFs=4.1,Gs=33dB
6G: NFL=20.6,GL=19.6dB
Standard Cascade vs Additive Phase Noise Design (1/f!)
Pout=+26.1dBm
Pout=+30.2dBm
Pout=+28.6dBm
tj=4.5 fS,p-p
tj=3.15 fS,p-p

## Slide 45
X1 - 0dB - X4 ZX60-233LPN+
(A1: large signal. A2: SAT)
X1 - 8dB – X4 ZX60-233LPN+
(A1:small signal
A2: small signal)
Series Case: 6GHz
Series Case: 6GHz
X1 – 8dB-X4 ZX60-233LPN+
X1-6dB-X1 ZX60-233LPN+
SAT
LS
SS
SS
SS
SS
SS
LS
Red: +10.6/+16.8dBm
Red: +4.5/+5.8dBm
Red: +4.5/+5.8dBm
Green: -0.89/+9.7dBm
Pout=+26.1dBm
Pout=+22.8dBm
Pout=+26.1dBm
Pout=+30.2dBm
tj=4.5 fS,p-p
tj=3.15 fS,p-p
tj=3.15 fS,p-p
tj=4.0 fS,p-p

## Slide 46
Series
Case: 2GHz
X1 – 0dB-X4 ZX60-233LPN+
X1-0dB-X1 ZX60-233LPN+
SS
LS
SAT
SAT
Red: +11.1/+17.3dBm
Red: +5.5/+20.7dBm
SS
LS
Red: +5.5/+15dBm
Series
Case: 2GHz
X1 – 0dB-X4 ZX60-233LPN+
LS
SAT
Red: +11.1/+17.3dBm
X1-6dB-X1 ZX60-233LPN+
Pout=+31.4dBm
Pout=+24.5dBm
Pout=+25.1dBm
Pout=+31.4dBm
tj=10.0 fS,p-p
tj=9.8 fS,p-p
tj=7.5 fS,p-p
tj=7.5 fS,p-p

## Slide 47
References
[1] “Microwave and Wireless Synthesizers, Theory and Design, 2nd Edition”, Ulrich L. Rohde, Enrico Rubiola, Jerry C. Whitaker, 2021 John Wiley & Sons, Inc. Published 2021 by John Wiley & Sons, Inc.
[2] “Phase Noise in RF and Microwave Amplifiers”, Rodolphe Boudot and Enrico Rubiola, IEEE Transactions on Ultrasonics, Ferroelectrics and Frequency Control, Vol. 59, No. 12, December 2012.
[3] “A Study of the Correlation Between High-Frequency Noise and Phase Noise in Low-Noise Silicon-Based Transistors”, Gilles Cibiel, Laurent Escotte and Oliver Llopis, IEEE Transactions on Microwave Theory and Techniques, Vol. 52, No. 1, January 2004.
[4] “Noise Figure vs. PM Noise Measurements: A Study at Microwave Frequencies”, A. Hati, D. A. Howe, F. L. Walls and D. Walker, Proceedings of the 2003 IEEE International Frequency Control Symposium and PDA Exhibition Jointly with the 17th European Frequency and Time Forum.
[5] “Additive Phase Noise in Linear and High-Efficiency X-Band Power Amplifiers”, Jason Breitbarth, Srdjan Pajic’, Narisi Wang and Zoya Popovic’, University of Colorado, Boulder, CO.
[6] “Study of PM Noise and Noise Figure in Low Noise Amplifiers Working under Small and Large Signal Conditions”, N. Garmendia, J. Portilla, Member, IEEE Electricity and Electronics Department, University of the Basque Country, Apdo. 644, 48080 Bilbao, Spain.
[7] “Studies on small- and large-signal noise in Solid-State Amplifiers”, J. Portilla and R. Jauregui Department of Electricity and Electronics University of the Basque Country, UPV/EHU Barrio Sarriena sin, 48940, Leioa, Spain joaquin.portilla@ehu.es; Rigoberto jauregui@ehu.es. N. Garmendia RF group ESS-Bilbao Ed. Cosimet, Landabarri, 2, 48940 Leioa, Spain ngarmendia@essbilbao.org.
[8] “Effect of the Input and Output Impedance on Amplifier Large-Signal Noise Performances”, Nagore Garmendia and Joaquín Portilla, Member, IEEE, IEEE TRANSACTIONS ON MICROWAVE THEORY AND TECHNIQUES, VOL. 59, NO. 6, JUNE 2011.
[9] “A collapse of the cross-spectral function in phase noise metrology”, C.W. Nelson, A. Hati and A. Howe, National Institute of Standards and Technology.
[10] “1/f Noise in Proton Irradiated SiGe HBTs”, Zhenrong Jin, Guofu Niu, John D. Cressler, Cheryl J. Marshall, Paul W. Marshall, Hak S. Kim, Robert Reed, and David L. Harame, Auburn University, NASA GSFC, Consultant, Jackson and Tull Chartered Engineers, IBM Microelectronics.
[11] “Flicker noise in AlGaAs/GaAs of high electron mobility heterostructure field-effect transistor at cryogenic temperature”, Souheil Mouetsi1,2, Foudil Zouach1 and Djamil Rechem1,2 . 1 Department of electrical engineering, faculty of sciences and applied sciences, University Larbi Ben M’hidi,Oum El Bouaghi, Algeria. 2 Laboratory of active components and materials, University Larbi Ben M’hidi, Oum El Bouaghi. Algeria
Low-Level Radio Frequency (LLRF) Workshop 2025
Amplifier Additive Phase Modulation Noise Characterization and Testing and Theoretical Discussion of White and Flicker Noise Processes in both Small and Large Signal Regions of Operation* Mini-Circuits
*Class A, AB Only
Id#107

## Slide 48
Understanding Additive Phase Noise in RF & Microwave Amplifiers - Part 1
https://m.youtube.com/watch?v=s09qfKjMLqo
Understanding Additive Phase Noise in RF & Microwave Amplifiers - Part 2
https://m.youtube.com/watch?v=ZWP0-BcWarc&feature=youtu.be
Understanding Additive Phase Noise in RF & Microwave Amplifiers - Part 3
https://www.youtube.com/watch?v=MK6iualRipM
Low-Level Radio Frequency (LLRF) Workshop 2025
Amplifier Additive Phase Modulation Noise Characterization and Testing and Theoretical Discussion of White and Flicker Noise Processes in both Small and Large Signal Regions of Operation* Mini-Circuits
*Class A, AB Only
Id#107